## Read me
Allowable total error sources — clinical chemistry and immunochemistry (cobas pro)
What is already filled in
All 63 analytes and 133 control levels from 1 May 2025 – 30 June 2026. The 10 manuscript-panel analytes come first, then the supplementary analytes. On the 'Sigma' sheet, CV, target and analyser offset (bias) are measured values: CV = pooled within-lot CV of the worst analyser; target = mean of the control lot in use at the end of the period; bias = largest analyser offset from that mean. These are the values the rule-design tool uses.
Turkey MoH 2026 percentage limits are filled in for the analytes listed in its Table 1 (comparison source only).
What you fill in
1. 'TEa input': the YELLOW cells — percentage limits for Labquality, CLIA, Rili-BÄK, Noklus, RCPA, China WS/T 403, SEKK, Belgium, Ricos and any other source, and CVI / CVG from the EFLM Biological Variation Database. The EFLM optimal, desirable and minimum tiers are calculated.
2. Absolute limits ('ref.' columns) are for reference only. They are defined around the EQA or proficiency-testing sample's target value and are not converted at the IQC concentration; only the % limit is used.
3. 'References': version, full citation and access date of each source.
4. 'For the tool': one TEa % per analyte and source, to enter in the tool's 'Allowable error sources' row.
Source priority
Biological variation (EFLM; Ricos as the older compilation) is the primary basis. Labquality, CLIA, Rili-BÄK, Noklus, RCPA, China WS/T 403-2024, SEKK Dmax, Belgium WIV-ISP and Turkey MoH 2026 are comparison sources: none binds this laboratory; they show how far the attainable sigma depends on the specification chosen.
Colours
Yellow fill = for you to fill in · Blue text = measured or fixed inputs · Black text = formulas · Green text = values pulled from another sheet
Rili-BÄK: the % is Column 5 of Table B 1-2 a (permissible relative deviation in the EQA); Column 3 (IQC single-value / RMSD limit) is recorded in the reference column. RCPA: the % applies above the stated concentration and an absolute limit below it; only the % is used, and the note says when a control level lies below that concentration (there the % is stricter than RCPA's own limit).
Sigma colours: red < 3 · orange 3–4 · light green 4–6 · dark green ≥ 6 (as in the tool).
Calculations (identical to the tool)
Biological variation (Fraser 1997): TEa = 1.65 × kI × CVI + kB × √(CVI² + CVG²); kI/kB = 0.25/0.125 (optimal), 0.50/0.25 (desirable), 0.75/0.375 (minimum).
Sigma = (TEa % − ¦bias %¦) / CV %.
Notes on particular analytes
PTH: in routine use from September 2025. Vitamin B12 (total) to 19 March 2026, Active B12 from 21 February 2026. Rheumatoid factor not run after 15 June 2026. C-peptide L2: lot 301804 excluded after review (mixed control levels); the 2026 lot is used. Bilirubin and CRP L3: third control (Bili Ped, CRP L3).

## TEa input
TEa input — fill the yellow cells
Percentage limits as numbers (10 = 10%). 'ref.' columns are recorded for reference only and are not used. A source left empty is ignored.
Analyte | Group | Section | Unit | Labquality |  | CLIA |  | Rili-BÄK |  | Noklus |  | RCPA |  | China WS/T 403-2024 |  | SEKK Dmax | Belgium WIV-ISP |  | Turkey MoH 2026 |  |  | EFLM biological variation |  |  |  |  | Ricos (BV) |  |  | Other |  |  | Note
|  |  |  | % | Absolute (ref.) | % | Absolute (ref.) | % | Absolute (ref.) | % | Absolute (ref.) | % | Absolute (ref.) | % | Absolute (ref.) | % | % | Goal type (ref.) | DKD % | DKD absolute as printed (ref.) | IKK max CV % (ref.) | CVI % | CVG % | Optimal % | Desirable % | Minimum % | Optimal % | Desirable % | Minimum % | Source name | % | Absolute (ref.)
Sodium | Panel | Clinical chemistry | mmol/L | 2 |  |  | ±4 mmol/L (absolute only) | 3 | EQA col. 5: ±5% |  |  | 2 | ±3 mmol/L ≤150 mmol/L | 4 |  | 5 | 3.4 | State of the art | 9 |  | 4 | 0.5 | 0.7 | 0.31378 | 0.62756 | 0.94134 |  | 0.73 |  |  |  |  | EFLM 95% CI: CVI 0.0–4.3, CVG 0.4–9.7; RCPA: L1 (137) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Creatinine | Panel | Clinical chemistry | µmol/L | 8 |  | 10 | ±0.2 mg/dL (greater) | 11.5 | EQA col. 5: ±20% |  |  | 8 | ±8 µmol/L ≤100 µmol/L | 12 | ±6 µmol/L at the low end; threshold misprinted in source ("≤5 µmol/L … >4 mmol/L") | 15 | 10.9 | State of the art | 20 | ± 0.4 mg/dL | 5 | 4.4 | 16.1 | 3.9013 | 7.8026 | 11.7039 |  | 8.87 |  |  |  |  | RCPA: L1 (99.1) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
ALT | Panel | Clinical chemistry | U/L | 12 |  | 15 | ±6 U/L (greater) | 11.5 | EQA col. 5: ±21% |  |  | 12 | ±5 U/L ≤40 U/L | 15 | ±6 U/L (≤40 U/L) | 15 | 13.3 | State of the art | 20 | ± 6 U/L | 6 | 12.4 | 29.8 | 9.1496 | 18.2992 | 27.4488 |  | 27.48
Glucose | Panel | Clinical chemistry | mmol/L | 6 |  | 8 | ±6 mg/dL (greater) | 5 | EQA col. 5: ±8% |  |  | 8 | ±0.4 mmol/L ≤5.0 mmol/L | 7 | ±0.21 mmol/L (≤3 mmol/L) | 9 | 6.3 | State of the art | 10 | ± 6 mg/dL | 5 | 4.9 | 8.1 | 3.2046 | 6.4092 | 9.6138 |  | 6.96
Calcium | Panel | Clinical chemistry | mmol/L | 3 |  |  | ±1.0 mg/dL (absolute only) | 6 | EQA col. 5: ±10% |  |  | 4 | ±0.10 mmol/L ≤2.50 mmol/L | 5 | ±0.1 mmol/L (≤2 mmol/L) | 10 | 5 | State of the art | 10 | ± 1.0 mg/dL | 5 | 1.8 | 2.7 | 1.1481 | 2.2962 | 3.4444 |  | 2.55 |  |  |  |  | RCPA: L1 (2.24) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
TSH | Panel | Immunochemistry | mIU/L |  |  | 20 | ±0.2 mIU/L (greater) | 13.5 | EQA col. 5: ±24% |  |  | 20 | ±0.1 mU/L ≤0.5 mU/L | 20 | ±0.1 mIU/L (≤0.5 mIU/L) | 14 |  |  | 25 | ± 0.2 mIU/L | 10 | 17.9 | 36.1 | 12.4205 | 24.841 | 37.2616 |  | 23.7
hs-Troponin T | Panel | Immunochemistry | ng/L |  |  | 30 | ±0.2 ng/mL = 200 ng/L (greater) | 20 | EQA col. 5: ±33% | 15 |  | 20 | ±10 ng/L ≤50 ng/L | [8 empty] | 11.4 | 30.9 | 8.8195 | 17.639 | 26.4584 | [6 empty] | RCPA: L1 (18.8) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Ferritin | Panel | Immunochemistry | µg/L | 12 |  | 20 |  | 13.5 | EQA col. 5: ±25% |  |  | 15 | ±4 µg/L ≤27 µg/L | 25 | ±6 µg/L (≤24 µg/L) | 24 |  |  | 25 | ± 5 µg/L | 10 | 12.9 | [5 empty] | 16.9 |  |  |  |  | EFLM: CVI 12.9%, CVG not available — BV TEa not computed
PSA | Panel | Immunochemistry | µg/L |  |  | 20 | ±0.2 ng/mL (greater) | 15.5 | EQA col. 5: ±25% |  |  | 8 | ±0.4 µg/L ≤5.0 µg/L | 25 | ±0.75 µg/L (≤3 µg/L) | 15 |  |  | 25 | ± 0.2 µg/L | 10 | 6.8 | 42 | 8.1234 | 16.2467 | 24.3701 |  | 33.6 |  |  |  |  | RCPA: L1 (4.05) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
25-OH vitamin D | Panel | Immunochemistry | nmol/L |  |  |  |  | 25 | no EQA limit (col. 5) |  |  | 15 | ±9 nmol/L ≤60 nmol/L | [5 empty] | 25 | ± 5.0 µg/L | 10 | 6.8 | 37.6 | 7.5812 | 15.1625 | 22.7437 |  |  |  | RfB | 30
Potassium | Supplementary | Clinical chemistry | mmol/L | 4 |  |  | ±0.3 mmol/L (absolute only) | 4.5 | EQA col. 5: ±8% |  |  | 5 | ±0.2 mmol/L ≤4.0 mmol/L | 6 | ±0.2 mmol/L (≤3.3 mmol/L) | 8 | 4.8 | Biological variability | 9 | ± 0.3 mmol/L | 4 | [6 empty] | 5.61 |  |  |  |  | RCPA: L1 (3.98) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Chloride | Supplementary | Clinical chemistry | mmol/L | 3 |  | 5 |  | 4.5 | EQA col. 5: ±8% |  |  | 3 | ±3 mmol/L ≤100 mmol/L | 4 |  | 7 | 5.1 | State of the art | [9 empty] | 1.5 |  |  |  |  | RCPA: L1 (97.6) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Urea | Supplementary | Clinical chemistry | mmol/L | 10 |  | 9 | ±2 mg/dL urea N (greater) | 10.5 | EQA col. 5: ±20% |  |  | 12 | ±0.5 mmol/L ≤4.0 mmol/L | 8 | ±0.32 mmol/L (≤4 mmol/L) | 15 | 9 | Biological variability | 20 | ± 4.3 mg/dL (urea) | 5 | [6 empty] | 15.55
Urate | Supplementary | Clinical chemistry | µmol/L | 8 |  | 10 |  | 7 | EQA col. 5: ±13% |  |  | 8 | ±0.030 mmol/L ≤0.380 mmol/L | 12 |  | 14 | 8 | State of the art | 13 | ± 0.5 mg/dL | 5 | [6 empty] | 11.97 |  |  |  |  | RCPA: L1 (289) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
AST | Supplementary | Clinical chemistry | U/L | 12 |  | 15 | ±6 U/L (greater) | 11.5 | EQA col. 5: ±21% |  |  | 12 | ±5 U/L ≤40 U/L | 15 | ±6 U/L (≤40 U/L) | 15 | 11.3 | Biological variability | 20 | ± 6 U/L | 6 | [6 empty] | 16.69
ALP | Supplementary | Clinical chemistry | U/L | 12 |  | 20 |  | 11 | EQA col. 5: ±18% |  |  | 12 | ±15 U/L ≤125 U/L | 18 | ±9 U/L (≤60 U/L) | 24 |  |  | 30 |  | 8 | [6 empty] | 12.04 |  |  |  |  | RCPA: L1 (90) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
GGT | Supplementary | Clinical chemistry | U/L | 12 |  | 15 | ±5 U/L (greater) | 11.5 | EQA col. 5: ±21% |  |  | 12 | ±5 U/L ≤40 U/L | 11 | ±4.4 U/L (≤40 U/L) | 15 | 14.2 | State of the art | 20 | ± 5 U/L | 6 | [6 empty] | 22.11
LD | Supplementary | Clinical chemistry | U/L | 12 |  | 15 |  | 9 | EQA col. 5: ±18% |  |  | 8 | ±20 U/L ≤250 U/L | 11 |  | 21 | 10.7 | State of the art | [9 empty] | 11.4 |  |  |  |  | RCPA: L1 (187) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
CK | Supplementary | Clinical chemistry | U/L | 12 |  | 20 |  | 11 | EQA col. 5: ±20% |  |  | 12 | ±15 U/L ≤125 U/L | 15 |  | 21 |  |  | 20 |  | 10 | [6 empty] | 30.3 |  |  |  |  | RCPA: L1 (101) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Amylase | Supplementary | Clinical chemistry | U/L | 12 |  | 20 |  | 7 | no EQA limit (col. 5) |  |  | 10 | ±10 U/L ≤100 U/L | 15 | ±9 U/L (≤50 U/L) | 15 | [11 empty] | 14.6
Lipase | Supplementary | Clinical chemistry | U/L | 12 |  |  |  | 11 | no EQA limit (col. 5) |  |  | 20 | ±12 U/L ≤60 U/L |  |  | 24 | 14.3 | State of the art | [9 empty] | 37.88 |  |  |  |  | RCPA: L1 (52.6) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Bilirubin | Supplementary | Clinical chemistry | µmol/L | 12 |  | 20 | ±0.4 mg/dL (greater) | 13 | EQA col. 5: ±22% |  |  | 12 | ±3 µmol/L ≤25 µmol/L | 15 | ±12.4 µmol/L (≤16 µmol/L) | 21 | 12 | State of the art | 20 | ± 0.4 mg/dL | 8 | [6 empty] | 26.94 |  |  |  |  | RCPA: L1 (17) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; Rili-BÄK col. 3: 13% applies >34 µmol/L (L2, L3); 22% at ≤34 µmol/L (L1 ≈ 17)
Albumin | Supplementary | Clinical chemistry | g/L | 5 |  | 8 |  | 12.5 | EQA col. 5: ±20% |  |  | 6 | ±2 g/L ≤33 g/L | 6 |  | 12 | 10.7 | State of the art | 15 |  | 6 | [6 empty] | 4.07 |  |  |  |  | RCPA: L1 (31.5) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Total protein | Supplementary | Clinical chemistry | g/L | 5 |  | 8 |  | 6 | EQA col. 5: ±10% |  |  | 5 | ±3 g/L ≤60 g/L | 5 |  | 9 | 6.8 | State of the art | 10 |  | 5 | [6 empty] | 3.63 |  |  |  |  | RCPA: L1 (51.7) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Phosphate | Supplementary | Clinical chemistry | mmol/L | 6 |  | 10 | ±0.3 mg/dL (greater) | 9 | EQA col. 5: ±16% |  |  | 8 | ±0.06 mmol/L ≤0.75 mmol/L | 10 |  | 15 | 7.4 | Biological variability | 16 | ± 0.3 mg/dL | 8 | [6 empty] | 10.11
Magnesium | Supplementary | Clinical chemistry | mmol/L | 6 |  | 15 |  | 7.5 | EQA col. 5: ±15% |  |  | 8 | ±0.10 mmol/L ≤1.25 mmol/L | 15 | ±0.12 mmol/L (≤0.8 mmol/L) | 15 | 8.9 | State of the art | 15 |  | 8 | [6 empty] | 4.8 |  |  |  |  | RCPA: L1 (0.797) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (0.797) at or below the cut-off (absolute limit there; % stricter)
CRP | Supplementary | Clinical chemistry | mg/L |  |  | 30 | ±1 mg/L (greater); hs-CRP | 13.5 | EQA col. 5: ±20% |  |  | 20 | ±0.8 mg/L ≤4.0 mg/L | [5 empty] | 20 |  | 8
Cholesterol | Supplementary | Clinical chemistry | mmol/L | 5 |  | 10 |  | 7 | EQA col. 5: ±13% |  |  | 6 | ±0.30 mmol/L ≤5.00 mmol/L | 9 |  | 8.5 | 6.5 | Biological variability | 13 |  | 5 | [6 empty] | 9.01 |  |  |  |  | RCPA: L1 (3.35) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
HDL cholesterol | Supplementary | Clinical chemistry | mmol/L | 10 |  | 20 | ±6 mg/dL (greater) | 13 | no EQA limit (col. 5) |  |  | 12 | ±0.10 mmol/L ≤0.80 mmol/L | 20 | ±0.16 mmol/L (≤0.8 mmol/L) | 15 | 15.3 | State of the art | 20 | ± 6 mg/dL | 8 | [6 empty] | 11.63
LDL cholesterol | Supplementary | Clinical chemistry | mmol/L | 10 |  | 20 | (direct) | 9 | no EQA limit (col. 5) |  |  | 10 | ±0.20 mmol/L ≤2.00 mmol/L | 20 | ±0.4 mmol/L (≤2 mmol/L) |  |  |  | 20 |  | 8 | [11 empty] | RCPA: L1 (1.85) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (1.85) at or below the cut-off (absolute limit there; % stricter)
Triglycerides | Supplementary | Clinical chemistry | mmol/L | 15 |  | 15 |  | 9 | EQA col. 5: ±16% |  |  | 12 | ±0.20 mmol/L ≤1.60 mmol/L | 14 |  | 15 | 20 | State of the art | 15 |  | 6 | [6 empty] | 25.99 |  |  |  |  | RCPA: L1 (0.939) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Iron | Supplementary | Clinical chemistry | µmol/L | 12 |  | 15 | [5 empty] | 12 | ±3 µmol/L ≤25 µmol/L | 15 | ±2.1 µmol/L (≤14 µmol/L) | 15 | 8.3 | State of the art | [9 empty] | 30.7 |  |  |  |  | RCPA: L1 (13.4) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (13.4) at or below the cut-off (absolute limit there; % stricter)
Transferrin | Supplementary | Clinical chemistry | g/L | 8 |  |  |  | 8 | EQA col. 5: ±12% |  |  | 8 | ±0.20 g/L ≤2.50 g/L | [19 empty] | RCPA: L1 (2.13) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
IgA | Supplementary | Clinical chemistry | g/L | 15 |  | 20 |  | 12 | EQA col. 5: ±20% |  |  |  |  | 20 |  | 20 | 10.6 | State of the art | [9 empty] | 13.5
IgG | Supplementary | Clinical chemistry | g/L | 8 |  | 20 |  | 10 | EQA col. 5: ±18% |  |  |  |  | 20 |  | 18 | 9.3 | State of the art | [9 empty] | 8
IgM | Supplementary | Clinical chemistry | g/L | 15 |  | 20 |  | 13 | EQA col. 5: ±26% |  |  |  |  | 25 |  | 21 | 12.2 | State of the art | [9 empty] | 16.8
Rheumatoid factor | Supplementary | Clinical chemistry | IU/mL
Bicarbonate | Supplementary | Clinical chemistry | mmol/L |  |  | 20 | (carbon dioxide) |  |  |  |  | 10 | ±2 mmol/L ≤20 mmol/L | [19 empty] | RCPA: L1 (18.8) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Ammonia | Supplementary | Clinical chemistry | µmol/L | [8 empty] | 10 | ±5 µmol/L ≤50 µmol/L
Ethanol | Supplementary | Clinical chemistry | mmol/L |  |  | 20 | (blood alcohol) | 9 | EQA col. 5: ±12% |  |  | 10 | ±2 mmol/L ≤20 mmol/L | [19 empty] | 21% at ≤0.6 g/L (L1 ≈ 0.52 g/L); RCPA: L1 (11.2) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; Rili-BÄK col. 3: 9% applies >0.6 g/L (L2); 15% at ≤0.6 g/L (L1 ≈ 0.52 g/L)
Homocysteine | Supplementary | Clinical chemistry | µmol/L | [8 empty] | 10 | ±1.5 µmol/L ≤15 µmol/L | [19 empty] | RCPA: L1 (12.7) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Lithium | Supplementary | Clinical chemistry | mmol/L |  |  | 15 | ±0.3 mmol/L (greater) | 6 | EQA col. 5: ±12% |  |  | 10 | ±0.20 mmol/L ≤2.0 mmol/L |  |  | 12 | [16 empty] | RCPA: L1 (0.552), L2 (1.59) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Gentamicin | Supplementary | Clinical chemistry | mg/L |  |  | 25 |  | 25 | no EQA limit (col. 5) |  |  | 10 | ±0.2 mg/L ≤2.0 mg/L | [19 empty] | RCPA: L1 (1.79) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Vancomycin | Supplementary | Clinical chemistry | mg/L |  |  | 15 | ±2 µg/mL (greater) | 12 | EQA col. 5: ±18% |  |  | 10 | ±2.0 mg/L ≤20.3 mg/L | [19 empty] | RCPA: L1 (7.02) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Valproate | Supplementary | Clinical chemistry | µmol/L |  |  | 20 |  | 11.5 | EQA col. 5: ±20% |  |  | 10 | ±4 mg/L ≤36 mg/L | [19 empty] | RCPA: L1 (234) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Carbamazepine | Supplementary | Clinical chemistry | µmol/L |  |  | 20 | ±1.0 µg/mL (greater) | 12 | EQA col. 5: ±20% |  |  | 10 | ±0.5 mg/L ≤4.7 mg/L | [19 empty] | RCPA: L1 (14.3) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Paracetamol | Supplementary | Clinical chemistry | µmol/L |  |  | 15 | ±3 µg/mL (greater) |  |  |  |  | 10 | ±3 mg/L ≤30 mg/L | [19 empty] | RCPA: L1 (97.1) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Free T4 | Supplementary | Immunochemistry | pmol/L |  |  | 15 | ±0.3 ng/dL (greater) | 13 | EQA col. 5: ±20% |  |  | 12 | ±1.5 pmol/L ≤12 pmol/L | 20 | ±2.4 pmol/L (≤12 pmol/L) | 13 |  |  | 20 | ± 0.4 ng/dL | 10 | [6 empty] | 8
Free T3 | Supplementary | Immunochemistry | pmol/L |  |  |  |  | 13 | EQA col. 5: ±20% |  |  | 20 | ±0.7 pmol/L ≤3.5 pmol/L | 20 | ±0.7 pmol/L (≤3.5 pmol/L)
NT-proBNP | Supplementary | Immunochemistry | ng/L |  |  | 30 |  | 15 | no EQA limit (col. 5) |  |  | 20 | ±25 ng/L ≤125 ng/L
PTH | Supplementary | Immunochemistry | pmol/L |  |  | 30 | [5 empty] | 12 | ±1.0 pmol/L ≤8.0 pmol/L | 30 | ±3 pmol/L (≤10 pmol/L) | [17 empty] | RCPA: L1 (1.64) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (1.64), L2 (8.35) at or below the cut-off (absolute limit there; % stricter)
Vitamin B12 (total) | Supplementary | Immunochemistry | pmol/L |  |  | 25 | ±30 pg/mL (greater) | 25 | no EQA limit (col. 5) |  |  | 15 | ±18 pmol/L ≤120 pmol/L | [5 empty] | 25 | ± 30 ng/L | 10
Active B12 | Supplementary | Immunochemistry | pmol/L
Folate | Supplementary | Immunochemistry | nmol/L |  |  | 30 | ±1 ng/mL (greater) | 25 | no EQA limit (col. 5) |  |  | 25 | ±1.5 nmol/L ≤6.0 nmol/L | 30 | ±2.4 nmol/L (≤8 nmol/L) | [17 empty] | RCPA: L1 (4.74) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (4.74) at or below the cut-off (absolute limit there; % stricter)
Procalcitonin | Supplementary | Immunochemistry | µg/L |  |  |  |  | 18 | no EQA limit (col. 5)
Cortisol | Supplementary | Immunochemistry | nmol/L |  |  | 20 |  | 16 | EQA col. 5: ±30% |  |  | 15 | ±15 nmol/L ≤100 nmol/L | 20 | ±20 nmol/L (≤100 nmol/L) | 16 | [11 empty] | 22.8 |  |  |  |  | Rili-BÄK col. 3: 16% applies >166 nmol/L (both levels)
hCG | Supplementary | Immunochemistry | IU/L |  |  | 18 | ±3 mIU/mL (greater) | 14 | EQA col. 5: ±30% |  |  | 10 | ±1 IU/L ≤10 IU/L | [5 empty] | 30 | ± 3 U/L | 10 | [11 empty] | RCPA: L1 (3.71) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; Rili-BÄK col. 3: 14% applies >100 IU/L (L2); 17% at ≤100 IU/L (L1 ≈ 3.7)
C-peptide | Supplementary | Immunochemistry | nmol/L | [8 empty] | 12 | ±0.15 nmol/L ≤1.25 nmol/L | 20 | ±0.25 nmol/L (≤1.25 nmol/L) | [17 empty] | RCPA: L1 (0.437), L2 (0.695) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (0.437), L2 (0.695) at or below the cut-off (absolute limit there; % stricter)
CEA | Supplementary | Immunochemistry | µg/L |  |  | 15 | ±1 ng/dL as printed (greater) | 14 | EQA col. 5: ±24% |  |  | 12 | ±0.6 µg/L ≤5.0 µg/L | 25 | ±1.5 µg/L (≤6 µg/L) | 16 | [11 empty] | 24.7 |  |  |  |  | RCPA: L1 (3.15) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (3.15) at or below the cut-off (absolute limit there; % stricter)
CA 125 | Supplementary | Immunochemistry | kU/L |  |  | 20 |  | 16 | no EQA limit (col. 5) |  |  | 12 | ±6 kU/L ≤50 kU/L | 25 | ±10 kIU/L (≤40 kIU/L) | [17 empty] | RCPA: L1 (29.7), L2 (46.3) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there; China: L1 (29.7) at or below the cut-off (absolute limit there; % stricter)
Anti-TPO | Supplementary | Immunochemistry | kIU/L
Anti-CCP | Supplementary | Immunochemistry | U/mL
Digoxin | Supplementary | Immunochemistry | nmol/L |  |  | 15 | ±0.2 ng/mL (greater) |  |  |  |  | 10 | ±0.2 µg/L ≤1.6 µg/L |  |  | 30 | [16 empty] | RCPA: L1 (1.14) at or below the RCPA cut-off, where RCPA uses an absolute limit; the % is stricter there
Note on G5: CLIA gives most limits as '±% or ±absolute, whichever is greater', around the PT sample target. Only the % limit is used.
Note on T6: Ek-1/A Table 1, row 20: DKD sapma sınırı HD ± %9; IKK CV ≤ 4%.
Note on T7: Ek-1/A Table 1, row 16: DKD sapma sınırı HD ± %20 or ± 0.4 mg/dL (whichever greater); IKK CV ≤ 5%.
Note on T8: Ek-1/A Table 1, row 6: DKD sapma sınırı HD ± %20 or ± 6 U/L (whichever greater); IKK CV ≤ 6%.
Note on T9: Ek-1/A Table 1, row 12: DKD sapma sınırı HD ± %10 or ± 6 mg/dL (whichever greater); IKK CV ≤ 5%.
Note on T10: Ek-1/A Table 1, row 14: DKD sapma sınırı HD ± %10 or ± 1.0 mg/dL (whichever greater); IKK CV ≤ 5%.
Note on T11: Ek-1/A Table 1, row 32: DKD sapma sınırı HD ± %25 or ± 0.2 mIU/L (whichever greater); IKK CV ≤ 10%.
Note on T13: Ek-1/A Table 1, row 29: DKD sapma sınırı HD ± %25 or ± 5 µg/L (whichever greater); IKK CV ≤ 10%.
Note on T14: Ek-1/A Table 1, row 33: DKD sapma sınırı HD ± %25 or ± 0.2 µg/L (whichever greater); IKK CV ≤ 10%.
Note on T15: Ek-1/A Table 1, row 36: DKD sapma sınırı HD ± %25 or ± 5.0 µg/L (whichever greater); IKK CV ≤ 10%.
Note on T16: Ek-1/A Table 1, row 19: DKD sapma sınırı HD ± %9 or ± 0.3 mmol/L (whichever greater); IKK CV ≤ 4%.
Note on T18: Ek-1/A Table 1, row 25: DKD sapma sınırı HD ± %20 or ± 4.3 mg/dL (urea) (whichever greater); IKK CV ≤ 5%.
Note on T19: Ek-1/A Table 1, row 26: DKD sapma sınırı HD ± %13 or ± 0.5 mg/dL (whichever greater); IKK CV ≤ 5%.
Note on T20: Ek-1/A Table 1, row 8: DKD sapma sınırı HD ± %20 or ± 6 U/L (whichever greater); IKK CV ≤ 6%.
Note on T21: Ek-1/A Table 1, row 7: DKD sapma sınırı HD ± %30; IKK CV ≤ 8%.
Note on T22: Ek-1/A Table 1, row 11: DKD sapma sınırı HD ± %20 or ± 5 U/L (whichever greater); IKK CV ≤ 6%.
Note on T24: Ek-1/A Table 1, row 15: DKD sapma sınırı HD ± %20; IKK CV ≤ 10%.
Note on T27: Ek-1/A Table 1, row 21: DKD sapma sınırı HD ± %20 or ± 0.4 mg/dL (whichever greater); IKK CV ≤ 8%.
Note on T28: Ek-1/A Table 1, row 5: DKD sapma sınırı HD ± %15; IKK CV ≤ 6%.
Note on T29: Ek-1/A Table 1, row 23: DKD sapma sınırı HD ± %10; IKK CV ≤ 5%.
Note on T30: Ek-1/A Table 1, row 10: DKD sapma sınırı HD ± %16 or ± 0.3 mg/dL (whichever greater); IKK CV ≤ 8%.
Note on T31: Ek-1/A Table 1, row 18: DKD sapma sınırı HD ± %15; IKK CV ≤ 8%.
Note on T32: Ek-1/A Table 1, row 9: DKD sapma sınırı HD ± %20; IKK CV ≤ 8%.
Note on T33: Ek-1/A Table 1, row 22: DKD sapma sınırı HD ± %13; IKK CV ≤ 5%.
Note on T34: Ek-1/A Table 1, row 13: DKD sapma sınırı HD ± %20 or ± 6 mg/dL (whichever greater); IKK CV ≤ 8%.
Note on T35: Ek-1/A Table 1, row 17: DKD sapma sınırı HD ± %20; IKK CV ≤ 8%.
Note on T36: Ek-1/A Table 1, row 24: DKD sapma sınırı HD ± %15; IKK CV ≤ 6%.
Note on T53: Ek-1/A Table 1, row 31: DKD sapma sınırı HD ± %20 or ± 0.4 ng/dL (whichever greater); IKK CV ≤ 10%.
Note on T57: Ek-1/A Table 1, row 35: DKD sapma sınırı HD ± %25 or ± 30 ng/L (whichever greater); IKK CV ≤ 10%.
Note on T62: Ek-1/A Table 1, row 30: DKD sapma sınırı HD ± %30 or ± 3 U/L (whichever greater); IKK CV ≤ 10%.
Note on W5: Within-subject CV (CVI) from the EFLM Biological Variation Database. Record the access date on the References sheet.

## Sigma
Effective TEa and sigma — per control level
CV, target and bias are measured (1 May 2025 – 30 June 2026). Effective TEa = the % limit of each source. Sigma = (TEa − ¦bias¦) / CV.
Analyte | Level | Control | Target (current lot mean) | CV % (worst analyser) | Bias % (largest analyser offset) | n | Worst analyser | Row | Effective TEa % | [15 empty] | Sigma
[9 empty] | Labquality | CLIA | Rili-BÄK | Noklus | RCPA | China WS/T 403-2024 | SEKK Dmax | Belgium WIV-ISP | Turkey MoH 2026 | BV EFLM (optimal) | BV EFLM (desirable) | BV EFLM (minimum) | BV Ricos (optimal) | BV Ricos (desirable) | BV Ricos (minimum) | Other | Labquality | CLIA | Rili-BÄK | Noklus | RCPA | China WS/T 403-2024 | SEKK Dmax | Belgium WIV-ISP | Turkey MoH 2026 | BV EFLM (optimal) | BV EFLM (desirable) | BV EFLM (minimum) | BV Ricos (optimal) | BV Ricos (desirable) | BV Ricos (minimum) | Other
Sodium | L1 | ClinChem 2 | 136.569 | 1.088 | 0.236 | 1777 | NISE-2.2 | 1 | 2 |  | 3 |  | 2 | 4 | 5 | 3.4 | 9 | 0.31378 | 0.62756 | 0.94134 |  | 0.73 |  |  | 1.6213 |  | 2.5404 |  | 1.6213 | 3.4596 | 4.3787 | 2.9081 | 8.0551 | 0.071488 | 0.35989 | 0.64829 |  | 0.45404
Sodium | L2 | ClinChem 3 | 154.677 | 0.726 | 0.251 | 1787 | NISE-2.2 | 1 | 2 |  | 3 |  | 2 | 4 | 5 | 3.4 | 9 | 0.31378 | 0.62756 | 0.94134 |  | 0.73 |  |  | 2.4091 |  | 3.7865 |  | 2.4091 | 5.1639 | 6.5413 | 4.3375 | 12.051 | 0.086473 | 0.51868 | 0.95088 |  | 0.65978
Creatinine | L1 | ClinChem 2 | 99.1432 | 1.903 | 0.116 | 1325 | c703-1.1 | 2 | 8 | 10 | 11.5 |  | 8 | 12 | 15 | 10.9 | 20 | 3.9013 | 7.8026 | 11.7039 |  | 8.87 |  |  | 4.1429 | 5.1939 | 5.9821 |  | 4.1429 | 6.2449 | 7.8213 | 5.6668 | 10.4488 | 1.9891 | 4.0392 | 6.0893 |  | 4.6001
Creatinine | L2 | ClinChem 3 | 356.985 | 1.288 | 0.073 | 1303 | c703-1.1 | 2 | 8 | 10 | 11.5 |  | 8 | 12 | 15 | 10.9 | 20 | 3.9013 | 7.8026 | 11.7039 |  | 8.87 |  |  | 6.1545 | 7.7073 | 8.8719 |  | 6.1545 | 9.2601 | 11.5893 | 8.4061 | 15.4713 | 2.9723 | 6.0012 | 9.0302 |  | 6.83
ALT | L1 | ClinChem 2 | 53.6702 | 2.446 | 0.935 | 1180 | c703-1.2 | 3 | 12 | 15 | 11.5 |  | 12 | 15 | 15 | 13.3 | 20 | 9.1496 | 18.2992 | 27.4488 |  | 27.48 |  |  | 4.5237 | 5.7502 | 4.3193 |  | 4.5237 | 5.7502 | 5.7502 | 5.0552 | 7.7944 | 3.3584 | 7.099 | 10.8397 |  | 10.8524
ALT | L2 | ClinChem 3 | 166.062 | 2.171 | 1.059 | 1161 | c703-1.2 | 3 | 12 | 15 | 11.5 |  | 12 | 15 | 15 | 13.3 | 20 | 9.1496 | 18.2992 | 27.4488 |  | 27.48 |  |  | 5.0396 | 6.4215 | 4.8093 |  | 5.0396 | 6.4215 | 6.4215 | 5.6384 | 8.7246 | 3.7267 | 7.9411 | 12.1556 |  | 12.17
Glucose | L1 | ClinChem 2 | 5.3043 | 1.932 | 0.478 | 882 | c703-2.1 | 4 | 6 | 8 | 5 |  | 8 | 7 | 9 | 6.3 | 10 | 3.2046 | 6.4092 | 9.6138 |  | 6.96 |  |  | 2.8582 | 3.8934 | 2.3406 |  | 3.8934 | 3.3758 | 4.411 | 3.0135 | 4.9286 | 1.4113 | 3.07 | 4.7287 |  | 3.3551
Glucose | L2 | ClinChem 3 | 14.679 | 1.495 | 0.322 | 881 | c703-1.1 | 4 | 6 | 8 | 5 |  | 8 | 7 | 9 | 6.3 | 10 | 3.2046 | 6.4092 | 9.6138 |  | 6.96 |  |  | 3.798 | 5.1358 | 3.1291 |  | 5.1358 | 4.4669 | 5.8047 | 3.9987 | 6.4736 | 1.9282 | 4.0717 | 6.2152 |  | 4.4401
Calcium | L1 | ClinChem 2 | 2.2443 | 1.19 | 0.67 | 897 | c703-1.1 | 5 | 3 |  | 6 |  | 4 | 5 | 10 | 5 | 10 | 1.1481 | 2.2962 | 3.4444 |  | 2.55 |  |  | 1.958 |  | 4.479 |  | 2.7983 | 3.6387 | 7.8403 | 3.6387 | 7.8403 | 0.40179 | 1.3666 | 2.3314 |  | 1.5798
Calcium | L2 | ClinChem 3 | 3.1305 | 0.998 | 0.476 | 889 | c703-1.1 | 5 | 3 |  | 6 |  | 4 | 5 | 10 | 5 | 10 | 1.1481 | 2.2962 | 3.4444 |  | 2.55 |  |  | 2.5291 |  | 5.5351 |  | 3.5311 | 4.5331 | 9.5431 | 4.5331 | 9.5431 | 0.67347 | 1.8239 | 2.9743 |  | 2.0782
TSH | L1 | IMMUN L1 | 3.0294 | 2.462 | 0.894 | 1866 | e801-2.1 | 6 |  | 20 | 13.5 |  | 20 | 20 | 14 |  | 25 | 12.4205 | 24.841 | 37.2616 |  | 23.7 |  |  |  | 7.7604 | 5.1202 |  | 7.7604 | 7.7604 | 5.3233 |  | 9.7912 | 4.6818 | 9.7267 | 14.7716 |  | 9.2632
TSH | L2 | IMMUN L2 | 10.4928 | 2.287 | 0.874 | 1856 | e801-1.1 | 6 |  | 20 | 13.5 |  | 20 | 20 | 14 |  | 25 | 12.4205 | 24.841 | 37.2616 |  | 23.7 |  |  |  | 8.3629 | 5.5208 |  | 8.3629 | 8.3629 | 5.7394 |  | 10.5492 | 5.0488 | 10.4797 | 15.9106 |  | 9.9808
hs-Troponin T | L1 | IMMUN L1 | 18.7954 | 3.013 | 0.821 | 1012 | e801-2.1 | 7 |  | 30 | 20 | 15 | 20 |  |  |  |  | 8.8195 | 17.639 | 26.4584 | [5 empty] | 9.6844 | 6.3654 | 4.7059 | 6.3654 |  |  |  |  | 2.6547 | 5.5818 | 8.5089
hs-Troponin T | L2 | IMMUN L2 | 2134.1 | 4.509 | 0.402 | 1007 | e801-1.1 | 7 |  | 30 | 20 | 15 | 20 |  |  |  |  | 8.8195 | 17.639 | 26.4584 | [5 empty] | 6.5642 | 4.3464 | 3.2375 | 4.3464 |  |  |  |  | 1.8668 | 3.8228 | 5.7788
Ferritin | L1 | IMMUN L1 | 93.418 | 2.571 | 0.499 | 1886 | e801-2.1 | 8 | 12 | 20 | 13.5 |  | 15 | 25 | 24 |  | 25 |  |  |  |  | 16.9 |  |  | 4.4734 | 7.585 | 5.0568 |  | 5.6402 | 9.5298 | 9.1408 |  | 9.5298 |  |  |  |  | 6.3792
Ferritin | L2 | IMMUN L2 | 586.639 | 2.878 | 0.144 | 1871 | e801-1.2 | 8 | 12 | 20 | 13.5 |  | 15 | 25 | 24 |  | 25 |  |  |  |  | 16.9 |  |  | 4.1195 | 6.8992 | 4.6407 |  | 5.1619 | 8.6366 | 8.2891 |  | 8.6366 |  |  |  |  | 5.8221
PSA | L1 | IMMUN L1 | 4.046 | 4.088 | 2.859 | 988 | e801-1.1 | 9 |  | 20 | 15.5 |  | 8 | 25 | 15 |  | 25 | 8.1234 | 16.2467 | 24.3701 |  | 33.6 |  |  |  | 4.193 | 3.0922 |  | 1.2576 | 5.4161 | 2.9699 |  | 5.4161 | 1.2878 | 3.2749 | 5.262 |  | 7.5198
PSA | L2 | IMMUN L2 | 20.8947 | 3.992 | 2.653 | 987 | e801-1.1 | 9 |  | 20 | 15.5 |  | 8 | 25 | 15 |  | 25 | 8.1234 | 16.2467 | 24.3701 |  | 33.6 |  |  |  | 4.3454 | 3.2182 |  | 1.3394 | 5.5979 | 3.0929 |  | 5.5979 | 1.3703 | 3.4052 | 5.4402 |  | 7.7523
25-OH vitamin D | L1 | IMMUN L1 | 97.6875 | 6.697 | 1.798 | 1837 | e801-2.1 | 10 |  |  | 25 |  | 15 |  |  |  | 25 | 7.5812 | 15.1625 | 22.7437 |  |  |  | 30 |  |  | 3.4645 |  | 1.9713 |  |  |  | 3.4645 | 0.86356 | 1.9956 | 3.1276 |  |  |  | 4.2111
25-OH vitamin D | L2 | IMMUN L2 | 170.797 | 5.749 | 0.797 | 1802 | e801-2.1 | 10 |  |  | 25 |  | 15 |  |  |  | 25 | 7.5812 | 15.1625 | 22.7437 |  |  |  | 30 |  |  | 4.2099 |  | 2.4705 |  |  |  | 4.2099 | 1.1801 | 2.4988 | 3.8175 |  |  |  | 5.0797
Potassium | L1 | ClinChem 2 | 3.9838 | 1.26 | 0.247 | 1808 | NISE-2.2 | 11 | 4 |  | 4.5 |  | 5 | 6 | 8 | 4.8 | 9 |  |  |  |  | 5.61 |  |  | 2.9786 |  | 3.3754 |  | 3.7722 | 4.5659 | 6.1532 | 3.6135 | 6.9468 |  |  |  |  | 4.2563
Potassium | L2 | ClinChem 3 | 5.9691 | 0.824 | 0.245 | 1805 | NISE-2.1 | 11 | 4 |  | 4.5 |  | 5 | 6 | 8 | 4.8 | 9 |  |  |  |  | 5.61 |  |  | 4.557 |  | 5.1638 |  | 5.7706 | 6.9842 | 9.4114 | 5.5279 | 10.625 |  |  |  |  | 6.5109
Chloride | L1 | ClinChem 2 | 97.5954 | 1.184 | 0.193 | 1792 | NISE-2.2 | 12 | 3 | 5 | 4.5 |  | 3 | 4 | 7 | 5.1 | [5 empty] | 1.5 |  |  | 2.3708 | 4.06 | 3.6377 |  | 2.3708 | 3.2154 | 5.7492 | 4.1444 | [5 empty] | 1.1039
Chloride | L2 | ClinChem 3 | 119.793 | 0.927 | 0.254 | 1796 | NISE-2.1 | 12 | 3 | 5 | 4.5 |  | 3 | 4 | 7 | 5.1 | [5 empty] | 1.5 |  |  | 2.9622 | 5.1197 | 4.5804 |  | 2.9622 | 4.041 | 7.2772 | 5.2276 | [5 empty] | 1.3441
Urea | L1 | ClinChem 2 | 6.3495 | 3.572 | 0.469 | 935 | c703-1.2 | 13 | 10 | 9 | 10.5 |  | 12 | 8 | 15 | 9 | 20 |  |  |  |  | 15.55 |  |  | 2.6683 | 2.3883 | 2.8082 |  | 3.2282 | 2.1083 | 4.068 | 2.3883 | 5.4678 |  |  |  |  | 4.222
Urea | L2 | ClinChem 3 | 20.2041 | 3.49 | 0.19 | 941 | c703-1.2 | 13 | 10 | 9 | 10.5 |  | 12 | 8 | 15 | 9 | 20 |  |  |  |  | 15.55 |  |  | 2.8109 | 2.5244 | 2.9542 |  | 3.384 | 2.2378 | 4.2436 | 2.5244 | 5.6762 |  |  |  |  | 4.4011
Urate | L1 | ClinChem 2 | 288.646 | 1.53 | 0.058 | 890 | c703-2.1 | 14 | 8 | 10 | 7 |  | 8 | 12 | 14 | 8 | 13 |  |  |  |  | 11.97 |  |  | 5.1908 | 6.498 | 4.5373 |  | 5.1908 | 7.8052 | 9.1124 | 5.1908 | 8.4588 |  |  |  |  | 7.7856
Urate | L2 | ClinChem 3 | 566.231 | 1.401 | 0.705 | 890 | c703-2.1 | 14 | 8 | 10 | 7 |  | 8 | 12 | 14 | 8 | 13 |  |  |  |  | 11.97 |  |  | 5.207 | 6.6345 | 4.4932 |  | 5.207 | 8.0621 | 9.4897 | 5.207 | 8.7759 |  |  |  |  | 8.0407
AST | L1 | ClinChem 2 | 47.2857 | 2.038 | 0.644 | 1038 | c703-2.1 | 15 | 12 | 15 | 11.5 |  | 12 | 15 | 15 | 11.3 | 20 |  |  |  |  | 16.69 |  |  | 5.5721 | 7.0442 | 5.3268 |  | 5.5721 | 7.0442 | 7.0442 | 5.2287 | 9.4975 |  |  |  |  | 7.8734
AST | L2 | ClinChem 3 | 233.286 | 1.922 | 0.408 | 1031 | c703-2.1 | 15 | 12 | 15 | 11.5 |  | 12 | 15 | 15 | 11.3 | 20 |  |  |  |  | 16.69 |  |  | 6.0312 | 7.5921 | 5.7711 |  | 6.0312 | 7.5921 | 7.5921 | 5.667 | 10.1935 |  |  |  |  | 8.4714
ALP | L1 | ClinChem 2 | 90.0234 | 2.387 | 0.193 | 993 | c703-2.2 | 16 | 12 | 20 | 11 |  | 12 | 18 | 24 |  | 30 |  |  |  |  | 12.04 |  |  | 4.9464 | 8.2979 | 4.5274 |  | 4.9464 | 7.46 | 9.9736 |  | 12.4872 |  |  |  |  | 4.9631
ALP | L2 | ClinChem 3 | 275.889 | 2.014 | 0.304 | 996 | c703-2.2 | 16 | 12 | 20 | 11 |  | 12 | 18 | 24 |  | 30 |  |  |  |  | 12.04 |  |  | 5.8073 | 9.7795 | 5.3108 |  | 5.8073 | 8.7865 | 11.7656 |  | 14.7448 |  |  |  |  | 5.8272
GGT | L1 | ClinChem 2 | 42.9146 | 2.407 | 0.217 | 1056 | c703-2.2 | 17 | 12 | 15 | 11.5 |  | 12 | 11 | 15 | 14.2 | 20 |  |  |  |  | 22.11 |  |  | 4.8953 | 6.1417 | 4.6876 |  | 4.8953 | 4.4799 | 6.1417 | 5.8093 | 8.2189 |  |  |  |  | 9.0956
GGT | L2 | ClinChem 3 | 176.17 | 1.741 | 0.294 | 1047 | c703-1.2 | 17 | 12 | 15 | 11.5 |  | 12 | 11 | 15 | 14.2 | 20 |  |  |  |  | 22.11 |  |  | 6.7237 | 8.4469 | 6.4365 |  | 6.7237 | 6.1493 | 8.4469 | 7.9874 | 11.3188 |  |  |  |  | 12.5307
LD | L1 | ClinChem 2 | 186.501 | 2.196 | 0.513 | 924 | c703-1.2 | 18 | 12 | 15 | 9 |  | 8 | 11 | 21 | 10.7 | [5 empty] | 11.4 |  |  | 5.2309 | 6.597 | 3.8648 |  | 3.4094 | 4.7755 | 9.3292 | 4.6389 | [5 empty] | 4.9577
LD | L2 | ClinChem 3 | 476.232 | 1.914 | 0.534 | 915 | c703-1.2 | 18 | 12 | 15 | 9 |  | 8 | 11 | 21 | 10.7 | [5 empty] | 11.4 |  |  | 5.9906 | 7.558 | 4.4232 |  | 3.9007 | 5.4681 | 10.6928 | 5.3114 | [5 empty] | 5.6771
CK | L1 | ClinChem 2 | 100.839 | 3.434 | 0.366 | 924 | c703-2.1 | 19 | 12 | 20 | 11 |  | 12 | 15 | 21 |  | 20 |  |  |  |  | 30.3 |  |  | 3.3879 | 5.7175 | 3.0967 |  | 3.3879 | 4.2615 | 6.0087 |  | 5.7175 |  |  |  |  | 8.7169
CK | L2 | ClinChem 3 | 371.25 | 2.482 | 0.596 | 917 | c703-1.1 | 19 | 12 | 20 | 11 |  | 12 | 15 | 21 |  | 20 |  |  |  |  | 30.3 |  |  | 4.5947 | 7.8179 | 4.1918 |  | 4.5947 | 5.8034 | 8.2208 |  | 7.8179 |  |  |  |  | 11.9678
Amylase | L1 | ClinChem 2 | 100.975 | 1.065 | 0.179 | 937 | c703-1.2 | 20 | 12 | 20 | 7 |  | 10 | 15 | 15 | [6 empty] | 14.6 |  |  | 11.0995 | 18.6113 | 6.4047 |  | 9.2216 | 13.9164 | 13.9164 | [6 empty] | 13.5408
Amylase | L2 | ClinChem 3 | 408.954 | 0.913 | 0.2 | 926 | c703-1.2 | 20 | 12 | 20 | 7 |  | 10 | 15 | 15 | [6 empty] | 14.6 |  |  | 12.9244 | 21.6867 | 7.448 |  | 10.7338 | 16.2103 | 16.2103 | [6 empty] | 15.7722
Lipase | L1 | ClinChem 2 | 52.6128 | 4.17 | 0.069 | 986 | c703-1.2 | 21 | 12 |  | 11 |  | 20 |  | 24 | 14.3 | [5 empty] | 37.88 |  |  | 2.8612 |  | 2.6213 |  | 4.7796 |  | 5.7388 | 3.4127 | [5 empty] | 9.0674
Lipase | L2 | ClinChem 3 | 152.686 | 5.298 | 0.215 | 992 | c703-2.2 | 21 | 12 |  | 11 |  | 20 |  | 24 | 14.3 | [5 empty] | 37.88 |  |  | 2.2244 |  | 2.0357 |  | 3.7344 |  | 4.4894 | 2.6586 | [5 empty] | 7.1093
Bilirubin | L1 | ClinChem 2 | 17.0339 | 8.389 | 1.968 | 986 | c703-2.1 | 22 | 12 | 20 | 13 |  | 12 | 15 | 21 | 12 | 20 |  |  |  |  | 26.94 |  |  | 1.1959 | 2.1495 | 1.3151 |  | 1.1959 | 1.5535 | 2.2687 | 1.1959 | 2.1495 |  |  |  |  | 2.9768
Bilirubin | L2 | ClinChem 3 | 96.7107 | 4.642 | 0.946 | 984 | c703-2.1 | 22 | 12 | 20 | 13 |  | 12 | 15 | 21 | 12 | 20 |  |  |  |  | 26.94 |  |  | 2.3813 | 4.1047 | 2.5967 |  | 2.3813 | 3.0276 | 4.3201 | 2.3813 | 4.1047 |  |  |  |  | 5.5997
Bilirubin | L3 | BILI PED | 334.13 | 3.867 | 1.103 | 920 | c703-2.1 | 22 | 12 | 20 | 13 |  | 12 | 15 | 21 | 12 | 20 |  |  |  |  | 26.94 |  |  | 2.8179 | 4.8867 | 3.0765 |  | 2.8179 | 3.5937 | 5.1453 | 2.8179 | 4.8867 |  |  |  |  | 6.6814
Albumin | L1 | ClinChem 2 | 31.4612 | 2.256 | 0.879 | 1075 | c703-2.2 | 23 | 5 | 8 | 12.5 |  | 6 | 6 | 12 | 10.7 | 15 |  |  |  |  | 4.07 |  |  | 1.8267 | 3.1565 | 5.1512 |  | 2.2699 | 2.2699 | 4.9295 | 4.3533 | 6.2593 |  |  |  |  | 1.4145
Albumin | L2 | ClinChem 3 | 49.5071 | 2.333 | 0.884 | 1074 | c703-2.2 | 23 | 5 | 8 | 12.5 |  | 6 | 6 | 12 | 10.7 | 15 |  |  |  |  | 4.07 |  |  | 1.7643 | 3.0502 | 4.979 |  | 2.1929 | 2.1929 | 4.7647 | 4.2075 | 6.0506 |  |  |  |  | 1.3656
Total protein | L1 | ClinChem 2 | 51.6588 | 1.971 | 0.1 | 866 | c703-1.1 | 24 | 5 | 8 | 6 |  | 5 | 5 | 9 | 6.8 | 10 |  |  |  |  | 3.63 |  |  | 2.486 | 4.0081 | 2.9934 |  | 2.486 | 2.486 | 4.5155 | 3.3993 | 5.0228 |  |  |  |  | 1.791
Total protein | L2 | ClinChem 3 | 81.7679 | 1.915 | 0.894 | 868 | c703-1.1 | 24 | 5 | 8 | 6 |  | 5 | 5 | 9 | 6.8 | 10 |  |  |  |  | 3.63 |  |  | 2.1441 | 3.7107 | 2.6663 |  | 2.1441 | 2.1441 | 4.2329 | 3.0841 | 4.7551 |  |  |  |  | 1.4287
Phosphate | L1 | ClinChem 2 | 1.3404 | 1.571 | 0.452 | 896 | c703-2.2 | 25 | 6 | 10 | 9 |  | 8 | 10 | 15 | 7.4 | 16 |  |  |  |  | 10.11 |  |  | 3.5315 | 6.0777 | 5.4411 |  | 4.8046 | 6.0777 | 9.2603 | 4.4227 | 9.8969 |  |  |  |  | 6.1477
Phosphate | L2 | ClinChem 3 | 2.6474 | 1.227 | 0.453 | 899 | c703-1.2 | 25 | 6 | 10 | 9 |  | 8 | 10 | 15 | 7.4 | 16 |  |  |  |  | 10.11 |  |  | 4.5208 | 7.7808 | 6.9658 |  | 6.1508 | 7.7808 | 11.8557 | 5.6618 | 12.6707 |  |  |  |  | 7.8704
Magnesium | L1 | ClinChem 2 | 0.79652 | 1.146 | 0.678 | 915 | c703-1.2 | 26 | 6 | 15 | 7.5 |  | 8 | 15 | 15 | 8.9 | 15 |  |  |  |  | 4.8 |  |  | 4.644 | 12.4974 | 5.9529 |  | 6.3892 | 12.4974 | 12.4974 | 7.1745 | 12.4974 |  |  |  |  | 3.5969
Magnesium | L2 | ClinChem 3 | 1.4416 | 1.075 | 0.73 | 916 | c703-1.2 | 26 | 6 | 15 | 7.5 |  | 8 | 15 | 15 | 8.9 | 15 |  |  |  |  | 4.8 |  |  | 4.9023 | 13.2744 | 6.2977 |  | 6.7628 | 13.2744 | 13.2744 | 7.6 | 13.2744 |  |  |  |  | 3.786
CRP | L1 | ClinChem 2 | 5.9171 | 1.949 | 0.889 | 1181 | c703-1.1 | 27 |  | 30 | 13.5 |  | 20 |  |  |  | 20 | [8 empty] | 14.9364 | 6.4705 |  | 9.8055 |  |  |  | 9.8055
CRP | L2 | ClinChem 3 | 49.8194 | 1.89 | 0.748 | 1180 | c703-1.1 | 27 |  | 30 | 13.5 |  | 20 |  |  |  | 20 | [8 empty] | 15.4772 | 6.7471 |  | 10.1862 |  |  |  | 10.1862
CRP | L3 | CRP L3 | 73.7571 | 2.875 | 0.089 | 1162 | c703-1.1 | 27 |  | 30 | 13.5 |  | 20 |  |  |  | 20 | [8 empty] | 10.4038 | 4.6647 |  | 6.9256 |  |  |  | 6.9256
Cholesterol | L1 | ClinChem 2 | 3.3472 | 1.244 | 0.251 | 900 | c703-2.2 | 28 | 5 | 10 | 7 |  | 6 | 9 | 8.5 | 6.5 | 13 |  |  |  |  | 9.01 |  |  | 3.8175 | 7.8368 | 5.4252 |  | 4.6214 | 7.033 | 6.631 | 5.0233 | 10.2484 |  |  |  |  | 7.041
Cholesterol | L2 | ClinChem 3 | 5.4807 | 1.18 | 0.245 | 899 | c703-2.2 | 28 | 5 | 10 | 7 |  | 6 | 9 | 8.5 | 6.5 | 13 |  |  |  |  | 9.01 |  |  | 4.0297 | 8.2669 | 5.7246 |  | 4.8771 | 7.4195 | 6.9958 | 5.3008 | 10.8093 |  |  |  |  | 7.428
HDL cholesterol | L1 | ClinChem 2 | 0.89639 | 2.492 | 0.432 | 961 | c703-2.2 | 29 | 10 | 20 | 13 |  | 12 | 20 | 15 | 15.3 | 20 |  |  |  |  | 11.63 |  |  | 3.8395 | 7.8523 | 5.0433 |  | 4.6421 | 7.8523 | 5.8459 | 5.9663 | 7.8523 |  |  |  |  | 4.4936
HDL cholesterol | L2 | ClinChem 3 | 1.3889 | 2.637 | 0.437 | 961 | c703-2.2 | 29 | 10 | 20 | 13 |  | 12 | 20 | 15 | 15.3 | 20 |  |  |  |  | 11.63 |  |  | 3.6265 | 7.4187 | 4.7641 |  | 4.3849 | 7.4187 | 5.5226 | 5.6363 | 7.4187 |  |  |  |  | 4.2446
LDL cholesterol | L1 | ClinChem 2 | 1.8482 | 1.958 | 1.061 | 1003 | c703-2.2 | 30 | 10 | 20 | 9 |  | 10 | 20 |  |  | 20 | [7 empty] | 4.5654 | 9.6726 | 4.0546 |  | 4.5654 | 9.6726 |  |  | 9.6726
LDL cholesterol | L2 | ClinChem 3 | 3.0624 | 1.777 | 1.189 | 1000 | c703-2.2 | 30 | 10 | 20 | 9 |  | 10 | 20 |  |  | 20 | [7 empty] | 4.9584 | 10.5858 | 4.3956 |  | 4.9584 | 10.5858 |  |  | 10.5858
Triglycerides | L1 | ClinChem 2 | 0.93882 | 1.828 | 0.233 | 964 | c703-1.1 | 31 | 15 | 15 | 9 |  | 12 | 14 | 15 | 20 | 15 |  |  |  |  | 25.99 |  |  | 8.0782 | 8.0782 | 4.796 |  | 6.4371 | 7.5312 | 8.0782 | 10.8135 | 8.0782 |  |  |  |  | 14.0903
Triglycerides | L2 | ClinChem 3 | 2.9077 | 1.357 | 1.288 | 964 | c703-1.1 | 31 | 15 | 15 | 9 |  | 12 | 14 | 15 | 20 | 15 |  |  |  |  | 25.99 |  |  | 10.1046 | 10.1046 | 5.6831 |  | 7.8939 | 9.3677 | 10.1046 | 13.7892 | 10.1046 |  |  |  |  | 18.2034
Iron | L1 | ClinChem 2 | 13.3947 | 2.194 | 0.152 | 989 | c703-1.1 | 32 | 12 | 15 |  |  | 12 | 15 | 15 | 8.3 | [5 empty] | 30.7 |  |  | 5.4002 | 6.7675 |  |  | 5.4002 | 6.7675 | 6.7675 | 3.7138 | [5 empty] | 13.9234
Iron | L2 | ClinChem 3 | 37.5562 | 1.932 | 0.246 | 999 | c703-1.1 | 32 | 12 | 15 |  |  | 12 | 15 | 15 | 8.3 | [5 empty] | 30.7 |  |  | 6.0839 | 7.6366 |  |  | 6.0839 | 7.6366 | 7.6366 | 4.1687 | [5 empty] | 15.7629
Transferrin | L1 | ClinChem 2 | 2.1254 | 2.062 | 1.414 | 1047 | c703-1.1 | 33 | 8 |  | 8 |  | 8 | [11 empty] | 3.194 |  | 3.194 |  | 3.194
Transferrin | L2 | ClinChem 3 | 3.4627 | 2.415 | 2.05 | 1043 | c703-1.1 | 33 | 8 |  | 8 |  | 8 | [11 empty] | 2.4638 |  | 2.4638 |  | 2.4638
IgA | L1 | ClinChem 2 | 1.5824 | 1.928 | 1.512 | 933 | c703-2.2 | 34 | 15 | 20 | 12 |  |  | 20 | 20 | 10.6 | [5 empty] | 13.5 |  |  | 6.9959 | 9.5892 | 5.4398 |  |  | 9.5892 | 9.5892 | 4.7137 | [5 empty] | 6.2178
IgA | L2 | ClinChem 3 | 2.6455 | 2.315 | 1.137 | 933 | c703-2.2 | 34 | 15 | 20 | 12 |  |  | 20 | 20 | 10.6 | [5 empty] | 13.5 |  |  | 5.9883 | 8.1482 | 4.6924 |  |  | 8.1482 | 8.1482 | 4.0877 | [5 empty] | 5.3404
IgG | L1 | ClinChem 2 | 8.4058 | 1.262 | 0.421 | 927 | c703-2.2 | 35 | 8 | 20 | 10 |  |  | 20 | 18 | 9.3 | [5 empty] | 8 |  |  | 6.0055 | 15.5143 | 7.5903 |  |  | 15.5143 | 13.9295 | 7.0357 | [5 empty] | 6.0055
IgG | L2 | ClinChem 3 | 13.6334 | 1.322 | 0.361 | 927 | c703-2.2 | 35 | 8 | 20 | 10 |  |  | 20 | 18 | 9.3 | [5 empty] | 8 |  |  | 5.7784 | 14.8555 | 7.2912 |  |  | 14.8555 | 13.3427 | 6.7617 | [5 empty] | 5.7784
IgM | L1 | ClinChem 2 | 0.71297 | 4.664 | 2.477 | 942 | c703-2.2 | 36 | 15 | 20 | 13 |  |  | 25 | 21 | 12.2 | [5 empty] | 16.8 |  |  | 2.685 | 3.7571 | 2.2562 |  |  | 4.8291 | 3.9715 | 2.0847 | [5 empty] | 3.071
IgM | L2 | ClinChem 3 | 1.166 | 3.583 | 2.071 | 944 | c703-2.2 | 36 | 15 | 20 | 13 |  |  | 25 | 21 | 12.2 | [5 empty] | 16.8 |  |  | 3.6084 | 5.0039 | 3.0502 |  |  | 6.3994 | 5.283 | 2.827 | [5 empty] | 4.1108
Rheumatoid factor | L1 | RF 1 | 20.005 | 3.369 | 0.511 | 802 | c703-2.1 | 37
Rheumatoid factor | L2 | RF 2 | 50.4552 | 2.596 | 0.544 | 817 | c703-2.1 | 37
Bicarbonate | L1 | AEC-N | 18.7928 | 2.752 | 0.189 | 931 | c703-2.1 | 38 |  | 20 |  |  | 10 | [12 empty] | 7.1988 |  |  | 3.565
Bicarbonate | L2 | AEC-A | 31.9152 | 2.992 | 0.512 | 931 | c703-1.1 | 38 |  | 20 |  |  | 10 | [12 empty] | 6.5134 |  |  | 3.1711
Ammonia | L1 | AEC-N | 71.2384 | 3.1 | 0.169 | 866 | c703-1.1 | 39 |  |  |  |  | 10 | [15 empty] | 3.1713
Ammonia | L2 | AEC-A | 238.245 | 2.071 | 0.064 | 870 | c703-2.1 | 39 |  |  |  |  | 10 | [15 empty] | 4.7977
Ethanol | L1 | AEC-N | 11.2 | 2.805 | 0.11 | 789 | c703-2.2 | 40 |  | 20 | 9 |  | 10 | [12 empty] | 7.0909 | 3.1693 |  | 3.5258
Ethanol | L2 | AEC-A | 33.4212 | 2.108 | 0.407 | 789 | c703-2.2 | 40 |  | 20 | 9 |  | 10 | [12 empty] | 9.2946 | 4.0764 |  | 4.5508
Homocysteine | L1 | HCY 1 | 12.7367 | 2.463 | 0.4 | 1113 | c703-2.2 | 41 |  |  |  |  | 10 | [15 empty] | 3.8977
Homocysteine | L2 | HCY 2 | 38.7577 | 2.817 | 0.364 | 1115 | c703-2.2 | 41 |  |  |  |  | 10 | [15 empty] | 3.4207
Lithium | L1 | ClinChem 2 | 0.5516 | 2.632 | 0.506 | 872 | c703-2.1 | 42 |  | 15 | 6 |  | 10 |  | 12 | [10 empty] | 5.5068 | 2.0874 |  | 3.6071 |  | 4.367
Lithium | L2 | ClinChem 3 | 1.5884 | 1.504 | 0.441 | 875 | c703-1.1 | 42 |  | 15 | 6 |  | 10 |  | 12 | [10 empty] | 9.6802 | 3.6961 |  | 6.3557 |  | 7.6855
Gentamicin | L1 | TDM 1 | 1.7941 | 4.527 | 0.383 | 870 | c703-2.2 | 43 |  | 25 | 25 |  | 10 | [12 empty] | 5.4378 | 5.4378 |  | 2.1244
Gentamicin | L2 | TDM 2 | 4.618 | 3.127 | 0.042 | 884 | c703-1.1 | 43 |  | 25 | 25 |  | 10 | [12 empty] | 7.9815 | 7.9815 |  | 3.1845
Gentamicin | L3 | TDM 3 | 7.4497 | 2.956 | 0.745 | 891 | c703-2.2 | 43 |  | 25 | 25 |  | 10 | [12 empty] | 8.2053 | 8.2053 |  | 3.1309
Vancomycin | L1 | TDM 1 | 7.0174 | 3.954 | 0.092 | 907 | c703-1.1 | 44 |  | 15 | 12 |  | 10 | [12 empty] | 3.7704 | 3.0116 |  | 2.5058
Vancomycin | L2 | TDM 2 | 20.581 | 2.589 | 0.483 | 913 | c703-1.1 | 44 |  | 15 | 12 |  | 10 | [12 empty] | 5.6072 | 4.4484 |  | 3.6759
Vancomycin | L3 | TDM 3 | 32.7686 | 2.561 | 1.53 | 913 | c703-1.1 | 44 |  | 15 | 12 |  | 10 | [12 empty] | 5.2597 | 4.0882 |  | 3.3073
Valproate | L1 | TDM 1 | 233.947 | 4.4 | 0.148 | 891 | c703-2.2 | 45 |  | 20 | 11.5 |  | 10 | [12 empty] | 4.5118 | 2.58 |  | 2.2391
Valproate | L2 | TDM 2 | 489.779 | 3.681 | 0.299 | 890 | c703-1.2 | 45 |  | 20 | 11.5 |  | 10 | [12 empty] | 5.3521 | 3.0429 |  | 2.6354
Valproate | L3 | TDM 3 | 765.734 | 3.812 | 1.061 | 890 | c703-1.2 | 45 |  | 20 | 11.5 |  | 10 | [12 empty] | 4.9683 | 2.7385 |  | 2.345
Carbamazepine | L1 | TDM 1 | 14.2807 | 3.293 | 0.255 | 881 | c703-1.2 | 46 |  | 20 | 12 |  | 10 | [12 empty] | 5.9961 | 3.5667 |  | 2.9593
Carbamazepine | L2 | TDM 2 | 38.6691 | 2.104 | 0.121 | 879 | c703-1.2 | 46 |  | 20 | 12 |  | 10 | [12 empty] | 9.4482 | 5.6459 |  | 4.6953
Carbamazepine | L3 | TDM 3 | 64.1529 | 2.62 | 0.078 | 879 | c703-2.2 | 46 |  | 20 | 12 |  | 10 | [12 empty] | 7.6038 | 4.5504 |  | 3.787
Paracetamol | L1 | TDM 1 | 97.1421 | 2.829 | 0.584 | 861 | c703-2.2 | 47 |  | 15 |  |  | 10 | [12 empty] | 5.0958 |  |  | 3.3284
Paracetamol | L2 | TDM 2 | 265.425 | 3.134 | 0.004 | 857 | c703-2.2 | 47 |  | 15 |  |  | 10 | [12 empty] | 4.7849 |  |  | 3.1895
Paracetamol | L3 | TDM 3 | 738.456 | 4.182 | 0.206 | 844 | c703-2.2 | 47 |  | 15 |  |  | 10 | [12 empty] | 3.5375 |  |  | 2.3419
Free T4 | L1 | IMMUN L1 | 17.0291 | 3.519 | 0.663 | 1801 | e801-1.2 | 48 |  | 15 | 13 |  | 12 | 20 | 13 |  | 20 |  |  |  |  | 8 |  |  |  | 4.0742 | 3.5058 |  | 3.2217 | 5.495 | 3.5058 |  | 5.495 |  |  |  |  | 2.085
Free T4 | L2 | IMMUN L2 | 33.4832 | 3.119 | 1.284 | 1789 | e801-1.2 | 48 |  | 15 | 13 |  | 12 | 20 | 13 |  | 20 |  |  |  |  | 8 |  |  |  | 4.3976 | 3.7563 |  | 3.4357 | 6.0006 | 3.7563 |  | 6.0006 |  |  |  |  | 2.1533
Free T3 | L1 | IMMUN L1 | 5.5811 | 4.223 | 1.094 | 1003 | e801-2.1 | 49 |  |  | 13 |  | 20 | 20 | [12 empty] | 2.8193 |  | 4.4769 | 4.4769
Free T3 | L2 | IMMUN L2 | 12.2229 | 3.233 | 0.567 | 1000 | e801-2.1 | 49 |  |  | 13 |  | 20 | 20 | [12 empty] | 3.8457 |  | 6.0108 | 6.0108
NT-proBNP | L1 | IMMUN L1 | 164.611 | 3.806 | 0.135 | 1013 | e801-2.1 | 50 |  | 30 | 15 |  | 20 | [12 empty] | 7.8468 | 3.9057 |  | 5.2194
NT-proBNP | L2 | IMMUN L2 | 2231.85 | 3.817 | 0.501 | 997 | e801-2.1 | 50 |  | 30 | 15 |  | 20 | [12 empty] | 7.7283 | 3.7985 |  | 5.1085
PTH | L1 | IMMUN L1 | 1.6404 | 4.712 | 1.384 | 669 | e801-2.1 | 51 |  | 30 |  |  | 12 | 30 | [11 empty] | 6.073 |  |  | 2.253 | 6.073
PTH | L2 | IMMUN L2 | 8.3548 | 3.595 | 0.406 | 670 | e801-1.1 | 51 |  | 30 |  |  | 12 | 30 | [11 empty] | 8.232 |  |  | 3.225 | 8.232
Vitamin B12 (total) | L1 | IMMUN L1 | 288.393 | 4.371 | 3.777 | 1336 | e801-1.2 | 52 |  | 25 | 25 |  | 15 |  |  |  | 25 | [8 empty] | 4.8554 | 4.8554 |  | 2.5676 |  |  |  | 4.8554
Vitamin B12 (total) | L2 | IMMUN L2 | 513.432 | 3.418 | 2.288 | 1338 | e801-1.2 | 52 |  | 25 | 25 |  | 15 |  |  |  | 25 | [8 empty] | 6.6448 | 6.6448 |  | 3.7191 |  |  |  | 6.6448
Active B12 | L1 | PC AcB12 1 | 24.6927 | 3.937 | 0.516 | 532 | e801-2.1 | 53
Active B12 | L2 | PC AcB12 2 | 52.8522 | 3.734 | 0.785 | 519 | e801-2.1 | 53
Folate | L1 | IMMUN L1 | 4.738 | 12.395 | 1.777 | 1818 | e801-1.2 | 54 |  | 30 | 25 |  | 25 | 30 | [11 empty] | 2.277 | 1.8736 |  | 1.8736 | 2.277
Folate | L2 | IMMUN L2 | 19.5068 | 8.041 | 1.882 | 1823 | e801-2.1 | 54 |  | 30 | 25 |  | 25 | 30 | [11 empty] | 3.4968 | 2.875 |  | 2.875 | 3.4968
Procalcitonin | L1 | PCT 1 | 0.52304 | 3.219 | 0.597 | 858 | e801-2.1 | 55 |  |  | 18 | [15 empty] | 5.4063
Procalcitonin | L2 | PCT 2 | 9.5906 | 3.073 | 0.104 | 859 | e801-2.1 | 55 |  |  | 18 | [15 empty] | 5.8236
Cortisol | L1 | IMMUN L1 | 261.528 | 2.921 | 0.373 | 945 | e801-1.1 | 56 |  | 20 | 16 |  | 15 | 20 | 16 | [6 empty] | 22.8 |  |  |  | 6.7193 | 5.3499 |  | 5.0075 | 6.7193 | 5.3499 | [6 empty] | 7.6779
Cortisol | L2 | IMMUN L2 | 568.753 | 3.182 | 0.474 | 945 | e801-2.1 | 56 |  | 20 | 16 |  | 15 | 20 | 16 | [6 empty] | 22.8 |  |  |  | 6.1364 | 4.8793 |  | 4.5651 | 6.1364 | 4.8793 | [6 empty] | 7.0163
hCG | L1 | IMMUN L1 | 3.7138 | 4.286 | 0.358 | 960 | e801-1.1 | 57 |  | 18 | 14 |  | 10 |  |  |  | 30 | [8 empty] | 4.1162 | 3.1829 |  | 2.2497 |  |  |  | 6.916
hCG | L2 | IMMUN L2 | 389.108 | 3.224 | 2.026 | 958 | e801-1.1 | 57 |  | 18 | 14 |  | 10 |  |  |  | 30 | [8 empty] | 4.9547 | 3.714 |  | 2.4733 |  |  |  | 8.6768
C-peptide | L1 | IMMUN L1 | 0.43673 | 5.022 | 0.601 | 1168 | e801-2.1 | 58 |  |  |  |  | 12 | 20 | [14 empty] | 2.2698 | 3.8628
C-peptide | L2 | IMMUN L2 | 0.69511 | 2.448 | 0.672 | 368 | e801-1.1 | 58 |  |  |  |  | 12 | 20 | [14 empty] | 4.6275 | 7.8954
CEA | L1 | IMMUN L1 | 3.1489 | 4.847 | 2.343 | 944 | e801-1.1 | 59 |  | 15 | 14 |  | 12 | 25 | 16 | [6 empty] | 24.7 |  |  |  | 2.6113 | 2.405 |  | 1.9924 | 4.6744 | 2.8176 | [6 empty] | 4.6125
CEA | L2 | IMMUN L2 | 23.0944 | 4.177 | 0.677 | 946 | e801-1.1 | 59 |  | 15 | 14 |  | 12 | 25 | 16 | [6 empty] | 24.7 |  |  |  | 3.429 | 3.1896 |  | 2.7108 | 5.8231 | 3.6684 | [6 empty] | 5.7513
CA 125 | L1 | IMMUN L1 | 29.6606 | 4.423 | 0.757 | 922 | e801-1.1 | 60 |  | 20 | 16 |  | 12 | 25 | [11 empty] | 4.3507 | 3.4463 |  | 2.5419 | 5.4811
CA 125 | L2 | IMMUN L2 | 46.2921 | 4.296 | 0.812 | 927 | e801-1.1 | 60 |  | 20 | 16 |  | 12 | 25 | [11 empty] | 4.4665 | 3.5354 |  | 2.6043 | 5.6304
Anti-TPO | L1 | IMMUN L1 | 12.6797 | 13.319 | 1.624 | 996 | e801-2.1 | 61
Anti-TPO | L2 | IMMUN L2 | 26.1563 | 9.453 | 0.475 | 988 | e801-1.1 | 61
Anti-CCP | L1 | ACCP1 | 16.7813 | 7.687 | 5.003 | 1077 | e801-1.1 | 62
Anti-CCP | L2 | ACCP2 | 95.2543 | 3.373 | 1.234 | 1093 | e801-2.1 | 62
Digoxin | L1 | IMMUN L1 | 1.1381 | 6.316 | 0.181 | 1039 | e801-1.1 | 63 |  | 15 |  |  | 10 |  | 30 | [10 empty] | 2.3463 |  |  | 1.5546 |  | 4.7212
Digoxin | L2 | IMMUN L2 | 2.676 | 3.856 | 1.084 | 1050 | e801-1.1 | 63 |  | 15 |  |  | 10 |  | 30 | [10 empty] | 3.6089 |  |  | 2.3122 |  | 7.499

## For the tool
TEa to enter in the tool — analyte × source
One TEa % per analyte and source. 'Lowest sigma' is the weakest control level for that source; filter to hide empty sources.
Analyte | Group | Source (name in the tool) | TEa % | Lowest sigma across levels
Sodium | Panel | Labquality | 2 | 1.6213
Sodium | Panel | CLIA
Sodium | Panel | Rili-BÄK | 3 | 2.5404
Sodium | Panel | Noklus
Sodium | Panel | RCPA | 2 | 1.6213
Sodium | Panel | China WS/T 403-2024 | 4 | 3.4596
Sodium | Panel | SEKK Dmax | 5 | 4.3787
Sodium | Panel | Belgium WIV-ISP | 3.4 | 2.9081
Sodium | Panel | Turkey MoH 2026 | 9 | 8.0551
Sodium | Panel | BV EFLM (optimal) | 0.31378 | 0.071488
Sodium | Panel | BV EFLM (desirable) | 0.62756 | 0.35989
Sodium | Panel | BV EFLM (minimum) | 0.94134 | 0.64829
Sodium | Panel | BV Ricos (optimal)
Sodium | Panel | BV Ricos (desirable) | 0.73 | 0.45404
Sodium | Panel | BV Ricos (minimum)
Sodium | Panel | Other
Creatinine | Panel | Labquality | 8 | 4.1429
Creatinine | Panel | CLIA | 10 | 5.1939
Creatinine | Panel | Rili-BÄK | 11.5 | 5.9821
Creatinine | Panel | Noklus
Creatinine | Panel | RCPA | 8 | 4.1429
Creatinine | Panel | China WS/T 403-2024 | 12 | 6.2449
Creatinine | Panel | SEKK Dmax | 15 | 7.8213
Creatinine | Panel | Belgium WIV-ISP | 10.9 | 5.6668
Creatinine | Panel | Turkey MoH 2026 | 20 | 10.4488
Creatinine | Panel | BV EFLM (optimal) | 3.9013 | 1.9891
Creatinine | Panel | BV EFLM (desirable) | 7.8026 | 4.0392
Creatinine | Panel | BV EFLM (minimum) | 11.7039 | 6.0893
Creatinine | Panel | BV Ricos (optimal)
Creatinine | Panel | BV Ricos (desirable) | 8.87 | 4.6001
Creatinine | Panel | BV Ricos (minimum)
Creatinine | Panel | Other
ALT | Panel | Labquality | 12 | 4.5237
ALT | Panel | CLIA | 15 | 5.7502
ALT | Panel | Rili-BÄK | 11.5 | 4.3193
ALT | Panel | Noklus
ALT | Panel | RCPA | 12 | 4.5237
ALT | Panel | China WS/T 403-2024 | 15 | 5.7502
ALT | Panel | SEKK Dmax | 15 | 5.7502
ALT | Panel | Belgium WIV-ISP | 13.3 | 5.0552
ALT | Panel | Turkey MoH 2026 | 20 | 7.7944
ALT | Panel | BV EFLM (optimal) | 9.1496 | 3.3584
ALT | Panel | BV EFLM (desirable) | 18.2992 | 7.099
ALT | Panel | BV EFLM (minimum) | 27.4488 | 10.8397
ALT | Panel | BV Ricos (optimal)
ALT | Panel | BV Ricos (desirable) | 27.48 | 10.8524
ALT | Panel | BV Ricos (minimum)
ALT | Panel | Other
Glucose | Panel | Labquality | 6 | 2.8582
Glucose | Panel | CLIA | 8 | 3.8934
Glucose | Panel | Rili-BÄK | 5 | 2.3406
Glucose | Panel | Noklus
Glucose | Panel | RCPA | 8 | 3.8934
Glucose | Panel | China WS/T 403-2024 | 7 | 3.3758
Glucose | Panel | SEKK Dmax | 9 | 4.411
Glucose | Panel | Belgium WIV-ISP | 6.3 | 3.0135
Glucose | Panel | Turkey MoH 2026 | 10 | 4.9286
Glucose | Panel | BV EFLM (optimal) | 3.2046 | 1.4113
Glucose | Panel | BV EFLM (desirable) | 6.4092 | 3.07
Glucose | Panel | BV EFLM (minimum) | 9.6138 | 4.7287
Glucose | Panel | BV Ricos (optimal)
Glucose | Panel | BV Ricos (desirable) | 6.96 | 3.3551
Glucose | Panel | BV Ricos (minimum)
Glucose | Panel | Other
Calcium | Panel | Labquality | 3 | 1.958
Calcium | Panel | CLIA
Calcium | Panel | Rili-BÄK | 6 | 4.479
Calcium | Panel | Noklus
Calcium | Panel | RCPA | 4 | 2.7983
Calcium | Panel | China WS/T 403-2024 | 5 | 3.6387
Calcium | Panel | SEKK Dmax | 10 | 7.8403
Calcium | Panel | Belgium WIV-ISP | 5 | 3.6387
Calcium | Panel | Turkey MoH 2026 | 10 | 7.8403
Calcium | Panel | BV EFLM (optimal) | 1.1481 | 0.40179
Calcium | Panel | BV EFLM (desirable) | 2.2962 | 1.3666
Calcium | Panel | BV EFLM (minimum) | 3.4444 | 2.3314
Calcium | Panel | BV Ricos (optimal)
Calcium | Panel | BV Ricos (desirable) | 2.55 | 1.5798
Calcium | Panel | BV Ricos (minimum)
Calcium | Panel | Other
TSH | Panel | Labquality
TSH | Panel | CLIA | 20 | 7.7604
TSH | Panel | Rili-BÄK | 13.5 | 5.1202
TSH | Panel | Noklus
TSH | Panel | RCPA | 20 | 7.7604
TSH | Panel | China WS/T 403-2024 | 20 | 7.7604
TSH | Panel | SEKK Dmax | 14 | 5.3233
TSH | Panel | Belgium WIV-ISP
TSH | Panel | Turkey MoH 2026 | 25 | 9.7912
TSH | Panel | BV EFLM (optimal) | 12.4205 | 4.6818
TSH | Panel | BV EFLM (desirable) | 24.841 | 9.7267
TSH | Panel | BV EFLM (minimum) | 37.2616 | 14.7716
TSH | Panel | BV Ricos (optimal)
TSH | Panel | BV Ricos (desirable) | 23.7 | 9.2632
TSH | Panel | BV Ricos (minimum)
TSH | Panel | Other
hs-Troponin T | Panel | Labquality
hs-Troponin T | Panel | CLIA | 30 | 6.5642
hs-Troponin T | Panel | Rili-BÄK | 20 | 4.3464
hs-Troponin T | Panel | Noklus | 15 | 3.2375
hs-Troponin T | Panel | RCPA | 20 | 4.3464
hs-Troponin T | Panel | China WS/T 403-2024
hs-Troponin T | Panel | SEKK Dmax
hs-Troponin T | Panel | Belgium WIV-ISP
hs-Troponin T | Panel | Turkey MoH 2026
hs-Troponin T | Panel | BV EFLM (optimal) | 8.8195 | 1.8668
hs-Troponin T | Panel | BV EFLM (desirable) | 17.639 | 3.8228
hs-Troponin T | Panel | BV EFLM (minimum) | 26.4584 | 5.7788
hs-Troponin T | Panel | BV Ricos (optimal)
hs-Troponin T | Panel | BV Ricos (desirable)
hs-Troponin T | Panel | BV Ricos (minimum)
hs-Troponin T | Panel | Other
Ferritin | Panel | Labquality | 12 | 4.1195
Ferritin | Panel | CLIA | 20 | 6.8992
Ferritin | Panel | Rili-BÄK | 13.5 | 4.6407
Ferritin | Panel | Noklus
Ferritin | Panel | RCPA | 15 | 5.1619
Ferritin | Panel | China WS/T 403-2024 | 25 | 8.6366
Ferritin | Panel | SEKK Dmax | 24 | 8.2891
Ferritin | Panel | Belgium WIV-ISP
Ferritin | Panel | Turkey MoH 2026 | 25 | 8.6366
Ferritin | Panel | BV EFLM (optimal)
Ferritin | Panel | BV EFLM (desirable)
Ferritin | Panel | BV EFLM (minimum)
Ferritin | Panel | BV Ricos (optimal)
Ferritin | Panel | BV Ricos (desirable) | 16.9 | 5.8221
Ferritin | Panel | BV Ricos (minimum)
Ferritin | Panel | Other
PSA | Panel | Labquality
PSA | Panel | CLIA | 20 | 4.193
PSA | Panel | Rili-BÄK | 15.5 | 3.0922
PSA | Panel | Noklus
PSA | Panel | RCPA | 8 | 1.2576
PSA | Panel | China WS/T 403-2024 | 25 | 5.4161
PSA | Panel | SEKK Dmax | 15 | 2.9699
PSA | Panel | Belgium WIV-ISP
PSA | Panel | Turkey MoH 2026 | 25 | 5.4161
PSA | Panel | BV EFLM (optimal) | 8.1234 | 1.2878
PSA | Panel | BV EFLM (desirable) | 16.2467 | 3.2749
PSA | Panel | BV EFLM (minimum) | 24.3701 | 5.262
PSA | Panel | BV Ricos (optimal)
PSA | Panel | BV Ricos (desirable) | 33.6 | 7.5198
PSA | Panel | BV Ricos (minimum)
PSA | Panel | Other
25-OH vitamin D | Panel | Labquality
25-OH vitamin D | Panel | CLIA
25-OH vitamin D | Panel | Rili-BÄK | 25 | 3.4645
25-OH vitamin D | Panel | Noklus
25-OH vitamin D | Panel | RCPA | 15 | 1.9713
25-OH vitamin D | Panel | China WS/T 403-2024
25-OH vitamin D | Panel | SEKK Dmax
25-OH vitamin D | Panel | Belgium WIV-ISP
25-OH vitamin D | Panel | Turkey MoH 2026 | 25 | 3.4645
25-OH vitamin D | Panel | BV EFLM (optimal) | 7.5812 | 0.86356
25-OH vitamin D | Panel | BV EFLM (desirable) | 15.1625 | 1.9956
25-OH vitamin D | Panel | BV EFLM (minimum) | 22.7437 | 3.1276
25-OH vitamin D | Panel | BV Ricos (optimal)
25-OH vitamin D | Panel | BV Ricos (desirable)
25-OH vitamin D | Panel | BV Ricos (minimum)
25-OH vitamin D | Panel | RfB | 30 | 4.2111
Potassium | Supplementary | Labquality | 4 | 2.9786
Potassium | Supplementary | CLIA
Potassium | Supplementary | Rili-BÄK | 4.5 | 3.3754
Potassium | Supplementary | Noklus
Potassium | Supplementary | RCPA | 5 | 3.7722
Potassium | Supplementary | China WS/T 403-2024 | 6 | 4.5659
Potassium | Supplementary | SEKK Dmax | 8 | 6.1532
Potassium | Supplementary | Belgium WIV-ISP | 4.8 | 3.6135
Potassium | Supplementary | Turkey MoH 2026 | 9 | 6.9468
Potassium | Supplementary | BV EFLM (optimal)
Potassium | Supplementary | BV EFLM (desirable)
Potassium | Supplementary | BV EFLM (minimum)
Potassium | Supplementary | BV Ricos (optimal)
Potassium | Supplementary | BV Ricos (desirable) | 5.61 | 4.2563
Potassium | Supplementary | BV Ricos (minimum)
Potassium | Supplementary | Other
Chloride | Supplementary | Labquality | 3 | 2.3708
Chloride | Supplementary | CLIA | 5 | 4.06
Chloride | Supplementary | Rili-BÄK | 4.5 | 3.6377
Chloride | Supplementary | Noklus
Chloride | Supplementary | RCPA | 3 | 2.3708
Chloride | Supplementary | China WS/T 403-2024 | 4 | 3.2154
Chloride | Supplementary | SEKK Dmax | 7 | 5.7492
Chloride | Supplementary | Belgium WIV-ISP | 5.1 | 4.1444
Chloride | Supplementary | Turkey MoH 2026
Chloride | Supplementary | BV EFLM (optimal)
Chloride | Supplementary | BV EFLM (desirable)
Chloride | Supplementary | BV EFLM (minimum)
Chloride | Supplementary | BV Ricos (optimal)
Chloride | Supplementary | BV Ricos (desirable) | 1.5 | 1.1039
Chloride | Supplementary | BV Ricos (minimum)
Chloride | Supplementary | Other
Urea | Supplementary | Labquality | 10 | 2.6683
Urea | Supplementary | CLIA | 9 | 2.3883
Urea | Supplementary | Rili-BÄK | 10.5 | 2.8082
Urea | Supplementary | Noklus
Urea | Supplementary | RCPA | 12 | 3.2282
Urea | Supplementary | China WS/T 403-2024 | 8 | 2.1083
Urea | Supplementary | SEKK Dmax | 15 | 4.068
Urea | Supplementary | Belgium WIV-ISP | 9 | 2.3883
Urea | Supplementary | Turkey MoH 2026 | 20 | 5.4678
Urea | Supplementary | BV EFLM (optimal)
Urea | Supplementary | BV EFLM (desirable)
Urea | Supplementary | BV EFLM (minimum)
Urea | Supplementary | BV Ricos (optimal)
Urea | Supplementary | BV Ricos (desirable) | 15.55 | 4.222
Urea | Supplementary | BV Ricos (minimum)
Urea | Supplementary | Other
Urate | Supplementary | Labquality | 8 | 5.1908
Urate | Supplementary | CLIA | 10 | 6.498
Urate | Supplementary | Rili-BÄK | 7 | 4.4932
Urate | Supplementary | Noklus
Urate | Supplementary | RCPA | 8 | 5.1908
Urate | Supplementary | China WS/T 403-2024 | 12 | 7.8052
Urate | Supplementary | SEKK Dmax | 14 | 9.1124
Urate | Supplementary | Belgium WIV-ISP | 8 | 5.1908
Urate | Supplementary | Turkey MoH 2026 | 13 | 8.4588
Urate | Supplementary | BV EFLM (optimal)
Urate | Supplementary | BV EFLM (desirable)
Urate | Supplementary | BV EFLM (minimum)
Urate | Supplementary | BV Ricos (optimal)
Urate | Supplementary | BV Ricos (desirable) | 11.97 | 7.7856
Urate | Supplementary | BV Ricos (minimum)
Urate | Supplementary | Other
AST | Supplementary | Labquality | 12 | 5.5721
AST | Supplementary | CLIA | 15 | 7.0442
AST | Supplementary | Rili-BÄK | 11.5 | 5.3268
AST | Supplementary | Noklus
AST | Supplementary | RCPA | 12 | 5.5721
AST | Supplementary | China WS/T 403-2024 | 15 | 7.0442
AST | Supplementary | SEKK Dmax | 15 | 7.0442
AST | Supplementary | Belgium WIV-ISP | 11.3 | 5.2287
AST | Supplementary | Turkey MoH 2026 | 20 | 9.4975
AST | Supplementary | BV EFLM (optimal)
AST | Supplementary | BV EFLM (desirable)
AST | Supplementary | BV EFLM (minimum)
AST | Supplementary | BV Ricos (optimal)
AST | Supplementary | BV Ricos (desirable) | 16.69 | 7.8734
AST | Supplementary | BV Ricos (minimum)
AST | Supplementary | Other
ALP | Supplementary | Labquality | 12 | 4.9464
ALP | Supplementary | CLIA | 20 | 8.2979
ALP | Supplementary | Rili-BÄK | 11 | 4.5274
ALP | Supplementary | Noklus
ALP | Supplementary | RCPA | 12 | 4.9464
ALP | Supplementary | China WS/T 403-2024 | 18 | 7.46
ALP | Supplementary | SEKK Dmax | 24 | 9.9736
ALP | Supplementary | Belgium WIV-ISP
ALP | Supplementary | Turkey MoH 2026 | 30 | 12.4872
ALP | Supplementary | BV EFLM (optimal)
ALP | Supplementary | BV EFLM (desirable)
ALP | Supplementary | BV EFLM (minimum)
ALP | Supplementary | BV Ricos (optimal)
ALP | Supplementary | BV Ricos (desirable) | 12.04 | 4.9631
ALP | Supplementary | BV Ricos (minimum)
ALP | Supplementary | Other
GGT | Supplementary | Labquality | 12 | 4.8953
GGT | Supplementary | CLIA | 15 | 6.1417
GGT | Supplementary | Rili-BÄK | 11.5 | 4.6876
GGT | Supplementary | Noklus
GGT | Supplementary | RCPA | 12 | 4.8953
GGT | Supplementary | China WS/T 403-2024 | 11 | 4.4799
GGT | Supplementary | SEKK Dmax | 15 | 6.1417
GGT | Supplementary | Belgium WIV-ISP | 14.2 | 5.8093
GGT | Supplementary | Turkey MoH 2026 | 20 | 8.2189
GGT | Supplementary | BV EFLM (optimal)
GGT | Supplementary | BV EFLM (desirable)
GGT | Supplementary | BV EFLM (minimum)
GGT | Supplementary | BV Ricos (optimal)
GGT | Supplementary | BV Ricos (desirable) | 22.11 | 9.0956
GGT | Supplementary | BV Ricos (minimum)
GGT | Supplementary | Other
LD | Supplementary | Labquality | 12 | 5.2309
LD | Supplementary | CLIA | 15 | 6.597
LD | Supplementary | Rili-BÄK | 9 | 3.8648
LD | Supplementary | Noklus
LD | Supplementary | RCPA | 8 | 3.4094
LD | Supplementary | China WS/T 403-2024 | 11 | 4.7755
LD | Supplementary | SEKK Dmax | 21 | 9.3292
LD | Supplementary | Belgium WIV-ISP | 10.7 | 4.6389
LD | Supplementary | Turkey MoH 2026
LD | Supplementary | BV EFLM (optimal)
LD | Supplementary | BV EFLM (desirable)
LD | Supplementary | BV EFLM (minimum)
LD | Supplementary | BV Ricos (optimal)
LD | Supplementary | BV Ricos (desirable) | 11.4 | 4.9577
LD | Supplementary | BV Ricos (minimum)
LD | Supplementary | Other
CK | Supplementary | Labquality | 12 | 3.3879
CK | Supplementary | CLIA | 20 | 5.7175
CK | Supplementary | Rili-BÄK | 11 | 3.0967
CK | Supplementary | Noklus
CK | Supplementary | RCPA | 12 | 3.3879
CK | Supplementary | China WS/T 403-2024 | 15 | 4.2615
CK | Supplementary | SEKK Dmax | 21 | 6.0087
CK | Supplementary | Belgium WIV-ISP
CK | Supplementary | Turkey MoH 2026 | 20 | 5.7175
CK | Supplementary | BV EFLM (optimal)
CK | Supplementary | BV EFLM (desirable)
CK | Supplementary | BV EFLM (minimum)
CK | Supplementary | BV Ricos (optimal)
CK | Supplementary | BV Ricos (desirable) | 30.3 | 8.7169
CK | Supplementary | BV Ricos (minimum)
CK | Supplementary | Other
Amylase | Supplementary | Labquality | 12 | 11.0995
Amylase | Supplementary | CLIA | 20 | 18.6113
Amylase | Supplementary | Rili-BÄK | 7 | 6.4047
Amylase | Supplementary | Noklus
Amylase | Supplementary | RCPA | 10 | 9.2216
Amylase | Supplementary | China WS/T 403-2024 | 15 | 13.9164
Amylase | Supplementary | SEKK Dmax | 15 | 13.9164
Amylase | Supplementary | Belgium WIV-ISP
Amylase | Supplementary | Turkey MoH 2026
Amylase | Supplementary | BV EFLM (optimal)
Amylase | Supplementary | BV EFLM (desirable)
Amylase | Supplementary | BV EFLM (minimum)
Amylase | Supplementary | BV Ricos (optimal)
Amylase | Supplementary | BV Ricos (desirable) | 14.6 | 13.5408
Amylase | Supplementary | BV Ricos (minimum)
Amylase | Supplementary | Other
Lipase | Supplementary | Labquality | 12 | 2.2244
Lipase | Supplementary | CLIA
Lipase | Supplementary | Rili-BÄK | 11 | 2.0357
Lipase | Supplementary | Noklus
Lipase | Supplementary | RCPA | 20 | 3.7344
Lipase | Supplementary | China WS/T 403-2024
Lipase | Supplementary | SEKK Dmax | 24 | 4.4894
Lipase | Supplementary | Belgium WIV-ISP | 14.3 | 2.6586
Lipase | Supplementary | Turkey MoH 2026
Lipase | Supplementary | BV EFLM (optimal)
Lipase | Supplementary | BV EFLM (desirable)
Lipase | Supplementary | BV EFLM (minimum)
Lipase | Supplementary | BV Ricos (optimal)
Lipase | Supplementary | BV Ricos (desirable) | 37.88 | 7.1093
Lipase | Supplementary | BV Ricos (minimum)
Lipase | Supplementary | Other
Bilirubin | Supplementary | Labquality | 12 | 1.1959
Bilirubin | Supplementary | CLIA | 20 | 2.1495
Bilirubin | Supplementary | Rili-BÄK | 13 | 1.3151
Bilirubin | Supplementary | Noklus
Bilirubin | Supplementary | RCPA | 12 | 1.1959
Bilirubin | Supplementary | China WS/T 403-2024 | 15 | 1.5535
Bilirubin | Supplementary | SEKK Dmax | 21 | 2.2687
Bilirubin | Supplementary | Belgium WIV-ISP | 12 | 1.1959
Bilirubin | Supplementary | Turkey MoH 2026 | 20 | 2.1495
Bilirubin | Supplementary | BV EFLM (optimal)
Bilirubin | Supplementary | BV EFLM (desirable)
Bilirubin | Supplementary | BV EFLM (minimum)
Bilirubin | Supplementary | BV Ricos (optimal)
Bilirubin | Supplementary | BV Ricos (desirable) | 26.94 | 2.9768
Bilirubin | Supplementary | BV Ricos (minimum)
Bilirubin | Supplementary | Other
Albumin | Supplementary | Labquality | 5 | 1.7643
Albumin | Supplementary | CLIA | 8 | 3.0502
Albumin | Supplementary | Rili-BÄK | 12.5 | 4.979
Albumin | Supplementary | Noklus
Albumin | Supplementary | RCPA | 6 | 2.1929
Albumin | Supplementary | China WS/T 403-2024 | 6 | 2.1929
Albumin | Supplementary | SEKK Dmax | 12 | 4.7647
Albumin | Supplementary | Belgium WIV-ISP | 10.7 | 4.2075
Albumin | Supplementary | Turkey MoH 2026 | 15 | 6.0506
Albumin | Supplementary | BV EFLM (optimal)
Albumin | Supplementary | BV EFLM (desirable)
Albumin | Supplementary | BV EFLM (minimum)
Albumin | Supplementary | BV Ricos (optimal)
Albumin | Supplementary | BV Ricos (desirable) | 4.07 | 1.3656
Albumin | Supplementary | BV Ricos (minimum)
Albumin | Supplementary | Other
Total protein | Supplementary | Labquality | 5 | 2.1441
Total protein | Supplementary | CLIA | 8 | 3.7107
Total protein | Supplementary | Rili-BÄK | 6 | 2.6663
Total protein | Supplementary | Noklus
Total protein | Supplementary | RCPA | 5 | 2.1441
Total protein | Supplementary | China WS/T 403-2024 | 5 | 2.1441
Total protein | Supplementary | SEKK Dmax | 9 | 4.2329
Total protein | Supplementary | Belgium WIV-ISP | 6.8 | 3.0841
Total protein | Supplementary | Turkey MoH 2026 | 10 | 4.7551
Total protein | Supplementary | BV EFLM (optimal)
Total protein | Supplementary | BV EFLM (desirable)
Total protein | Supplementary | BV EFLM (minimum)
Total protein | Supplementary | BV Ricos (optimal)
Total protein | Supplementary | BV Ricos (desirable) | 3.63 | 1.4287
Total protein | Supplementary | BV Ricos (minimum)
Total protein | Supplementary | Other
Phosphate | Supplementary | Labquality | 6 | 3.5315
Phosphate | Supplementary | CLIA | 10 | 6.0777
Phosphate | Supplementary | Rili-BÄK | 9 | 5.4411
Phosphate | Supplementary | Noklus
Phosphate | Supplementary | RCPA | 8 | 4.8046
Phosphate | Supplementary | China WS/T 403-2024 | 10 | 6.0777
Phosphate | Supplementary | SEKK Dmax | 15 | 9.2603
Phosphate | Supplementary | Belgium WIV-ISP | 7.4 | 4.4227
Phosphate | Supplementary | Turkey MoH 2026 | 16 | 9.8969
Phosphate | Supplementary | BV EFLM (optimal)
Phosphate | Supplementary | BV EFLM (desirable)
Phosphate | Supplementary | BV EFLM (minimum)
Phosphate | Supplementary | BV Ricos (optimal)
Phosphate | Supplementary | BV Ricos (desirable) | 10.11 | 6.1477
Phosphate | Supplementary | BV Ricos (minimum)
Phosphate | Supplementary | Other
Magnesium | Supplementary | Labquality | 6 | 4.644
Magnesium | Supplementary | CLIA | 15 | 12.4974
Magnesium | Supplementary | Rili-BÄK | 7.5 | 5.9529
Magnesium | Supplementary | Noklus
Magnesium | Supplementary | RCPA | 8 | 6.3892
Magnesium | Supplementary | China WS/T 403-2024 | 15 | 12.4974
Magnesium | Supplementary | SEKK Dmax | 15 | 12.4974
Magnesium | Supplementary | Belgium WIV-ISP | 8.9 | 7.1745
Magnesium | Supplementary | Turkey MoH 2026 | 15 | 12.4974
Magnesium | Supplementary | BV EFLM (optimal)
Magnesium | Supplementary | BV EFLM (desirable)
Magnesium | Supplementary | BV EFLM (minimum)
Magnesium | Supplementary | BV Ricos (optimal)
Magnesium | Supplementary | BV Ricos (desirable) | 4.8 | 3.5969
Magnesium | Supplementary | BV Ricos (minimum)
Magnesium | Supplementary | Other
CRP | Supplementary | Labquality
CRP | Supplementary | CLIA | 30 | 10.4038
CRP | Supplementary | Rili-BÄK | 13.5 | 4.6647
CRP | Supplementary | Noklus
CRP | Supplementary | RCPA | 20 | 6.9256
CRP | Supplementary | China WS/T 403-2024
CRP | Supplementary | SEKK Dmax
CRP | Supplementary | Belgium WIV-ISP
CRP | Supplementary | Turkey MoH 2026 | 20 | 6.9256
CRP | Supplementary | BV EFLM (optimal)
CRP | Supplementary | BV EFLM (desirable)
CRP | Supplementary | BV EFLM (minimum)
CRP | Supplementary | BV Ricos (optimal)
CRP | Supplementary | BV Ricos (desirable)
CRP | Supplementary | BV Ricos (minimum)
CRP | Supplementary | Other
Cholesterol | Supplementary | Labquality | 5 | 3.8175
Cholesterol | Supplementary | CLIA | 10 | 7.8368
Cholesterol | Supplementary | Rili-BÄK | 7 | 5.4252
Cholesterol | Supplementary | Noklus
Cholesterol | Supplementary | RCPA | 6 | 4.6214
Cholesterol | Supplementary | China WS/T 403-2024 | 9 | 7.033
Cholesterol | Supplementary | SEKK Dmax | 8.5 | 6.631
Cholesterol | Supplementary | Belgium WIV-ISP | 6.5 | 5.0233
Cholesterol | Supplementary | Turkey MoH 2026 | 13 | 10.2484
Cholesterol | Supplementary | BV EFLM (optimal)
Cholesterol | Supplementary | BV EFLM (desirable)
Cholesterol | Supplementary | BV EFLM (minimum)
Cholesterol | Supplementary | BV Ricos (optimal)
Cholesterol | Supplementary | BV Ricos (desirable) | 9.01 | 7.041
Cholesterol | Supplementary | BV Ricos (minimum)
Cholesterol | Supplementary | Other
HDL cholesterol | Supplementary | Labquality | 10 | 3.6265
HDL cholesterol | Supplementary | CLIA | 20 | 7.4187
HDL cholesterol | Supplementary | Rili-BÄK | 13 | 4.7641
HDL cholesterol | Supplementary | Noklus
HDL cholesterol | Supplementary | RCPA | 12 | 4.3849
HDL cholesterol | Supplementary | China WS/T 403-2024 | 20 | 7.4187
HDL cholesterol | Supplementary | SEKK Dmax | 15 | 5.5226
HDL cholesterol | Supplementary | Belgium WIV-ISP | 15.3 | 5.6363
HDL cholesterol | Supplementary | Turkey MoH 2026 | 20 | 7.4187
HDL cholesterol | Supplementary | BV EFLM (optimal)
HDL cholesterol | Supplementary | BV EFLM (desirable)
HDL cholesterol | Supplementary | BV EFLM (minimum)
HDL cholesterol | Supplementary | BV Ricos (optimal)
HDL cholesterol | Supplementary | BV Ricos (desirable) | 11.63 | 4.2446
HDL cholesterol | Supplementary | BV Ricos (minimum)
HDL cholesterol | Supplementary | Other
LDL cholesterol | Supplementary | Labquality | 10 | 4.5654
LDL cholesterol | Supplementary | CLIA | 20 | 9.6726
LDL cholesterol | Supplementary | Rili-BÄK | 9 | 4.0546
LDL cholesterol | Supplementary | Noklus
LDL cholesterol | Supplementary | RCPA | 10 | 4.5654
LDL cholesterol | Supplementary | China WS/T 403-2024 | 20 | 9.6726
LDL cholesterol | Supplementary | SEKK Dmax
LDL cholesterol | Supplementary | Belgium WIV-ISP
LDL cholesterol | Supplementary | Turkey MoH 2026 | 20 | 9.6726
LDL cholesterol | Supplementary | BV EFLM (optimal)
LDL cholesterol | Supplementary | BV EFLM (desirable)
LDL cholesterol | Supplementary | BV EFLM (minimum)
LDL cholesterol | Supplementary | BV Ricos (optimal)
LDL cholesterol | Supplementary | BV Ricos (desirable)
LDL cholesterol | Supplementary | BV Ricos (minimum)
LDL cholesterol | Supplementary | Other
Triglycerides | Supplementary | Labquality | 15 | 8.0782
Triglycerides | Supplementary | CLIA | 15 | 8.0782
Triglycerides | Supplementary | Rili-BÄK | 9 | 4.796
Triglycerides | Supplementary | Noklus
Triglycerides | Supplementary | RCPA | 12 | 6.4371
Triglycerides | Supplementary | China WS/T 403-2024 | 14 | 7.5312
Triglycerides | Supplementary | SEKK Dmax | 15 | 8.0782
Triglycerides | Supplementary | Belgium WIV-ISP | 20 | 10.8135
Triglycerides | Supplementary | Turkey MoH 2026 | 15 | 8.0782
Triglycerides | Supplementary | BV EFLM (optimal)
Triglycerides | Supplementary | BV EFLM (desirable)
Triglycerides | Supplementary | BV EFLM (minimum)
Triglycerides | Supplementary | BV Ricos (optimal)
Triglycerides | Supplementary | BV Ricos (desirable) | 25.99 | 14.0903
Triglycerides | Supplementary | BV Ricos (minimum)
Triglycerides | Supplementary | Other
Iron | Supplementary | Labquality | 12 | 5.4002
Iron | Supplementary | CLIA | 15 | 6.7675
Iron | Supplementary | Rili-BÄK
Iron | Supplementary | Noklus
Iron | Supplementary | RCPA | 12 | 5.4002
Iron | Supplementary | China WS/T 403-2024 | 15 | 6.7675
Iron | Supplementary | SEKK Dmax | 15 | 6.7675
Iron | Supplementary | Belgium WIV-ISP | 8.3 | 3.7138
Iron | Supplementary | Turkey MoH 2026
Iron | Supplementary | BV EFLM (optimal)
Iron | Supplementary | BV EFLM (desirable)
Iron | Supplementary | BV EFLM (minimum)
Iron | Supplementary | BV Ricos (optimal)
Iron | Supplementary | BV Ricos (desirable) | 30.7 | 13.9234
Iron | Supplementary | BV Ricos (minimum)
Iron | Supplementary | Other
Transferrin | Supplementary | Labquality | 8 | 2.4638
Transferrin | Supplementary | CLIA
Transferrin | Supplementary | Rili-BÄK | 8 | 2.4638
Transferrin | Supplementary | Noklus
Transferrin | Supplementary | RCPA | 8 | 2.4638
Transferrin | Supplementary | China WS/T 403-2024
Transferrin | Supplementary | SEKK Dmax
Transferrin | Supplementary | Belgium WIV-ISP
Transferrin | Supplementary | Turkey MoH 2026
Transferrin | Supplementary | BV EFLM (optimal)
Transferrin | Supplementary | BV EFLM (desirable)
Transferrin | Supplementary | BV EFLM (minimum)
Transferrin | Supplementary | BV Ricos (optimal)
Transferrin | Supplementary | BV Ricos (desirable)
Transferrin | Supplementary | BV Ricos (minimum)
Transferrin | Supplementary | Other
IgA | Supplementary | Labquality | 15 | 5.9883
IgA | Supplementary | CLIA | 20 | 8.1482
IgA | Supplementary | Rili-BÄK | 12 | 4.6924
IgA | Supplementary | Noklus
IgA | Supplementary | RCPA
IgA | Supplementary | China WS/T 403-2024 | 20 | 8.1482
IgA | Supplementary | SEKK Dmax | 20 | 8.1482
IgA | Supplementary | Belgium WIV-ISP | 10.6 | 4.0877
IgA | Supplementary | Turkey MoH 2026
IgA | Supplementary | BV EFLM (optimal)
IgA | Supplementary | BV EFLM (desirable)
IgA | Supplementary | BV EFLM (minimum)
IgA | Supplementary | BV Ricos (optimal)
IgA | Supplementary | BV Ricos (desirable) | 13.5 | 5.3404
IgA | Supplementary | BV Ricos (minimum)
IgA | Supplementary | Other
IgG | Supplementary | Labquality | 8 | 5.7784
IgG | Supplementary | CLIA | 20 | 14.8555
IgG | Supplementary | Rili-BÄK | 10 | 7.2912
IgG | Supplementary | Noklus
IgG | Supplementary | RCPA
IgG | Supplementary | China WS/T 403-2024 | 20 | 14.8555
IgG | Supplementary | SEKK Dmax | 18 | 13.3427
IgG | Supplementary | Belgium WIV-ISP | 9.3 | 6.7617
IgG | Supplementary | Turkey MoH 2026
IgG | Supplementary | BV EFLM (optimal)
IgG | Supplementary | BV EFLM (desirable)
IgG | Supplementary | BV EFLM (minimum)
IgG | Supplementary | BV Ricos (optimal)
IgG | Supplementary | BV Ricos (desirable) | 8 | 5.7784
IgG | Supplementary | BV Ricos (minimum)
IgG | Supplementary | Other
IgM | Supplementary | Labquality | 15 | 2.685
IgM | Supplementary | CLIA | 20 | 3.7571
IgM | Supplementary | Rili-BÄK | 13 | 2.2562
IgM | Supplementary | Noklus
IgM | Supplementary | RCPA
IgM | Supplementary | China WS/T 403-2024 | 25 | 4.8291
IgM | Supplementary | SEKK Dmax | 21 | 3.9715
IgM | Supplementary | Belgium WIV-ISP | 12.2 | 2.0847
IgM | Supplementary | Turkey MoH 2026
IgM | Supplementary | BV EFLM (optimal)
IgM | Supplementary | BV EFLM (desirable)
IgM | Supplementary | BV EFLM (minimum)
IgM | Supplementary | BV Ricos (optimal)
IgM | Supplementary | BV Ricos (desirable) | 16.8 | 3.071
IgM | Supplementary | BV Ricos (minimum)
IgM | Supplementary | Other
Rheumatoid factor | Supplementary | Labquality
Rheumatoid factor | Supplementary | CLIA
Rheumatoid factor | Supplementary | Rili-BÄK
Rheumatoid factor | Supplementary | Noklus
Rheumatoid factor | Supplementary | RCPA
Rheumatoid factor | Supplementary | China WS/T 403-2024
Rheumatoid factor | Supplementary | SEKK Dmax
Rheumatoid factor | Supplementary | Belgium WIV-ISP
Rheumatoid factor | Supplementary | Turkey MoH 2026
Rheumatoid factor | Supplementary | BV EFLM (optimal)
Rheumatoid factor | Supplementary | BV EFLM (desirable)
Rheumatoid factor | Supplementary | BV EFLM (minimum)
Rheumatoid factor | Supplementary | BV Ricos (optimal)
Rheumatoid factor | Supplementary | BV Ricos (desirable)
Rheumatoid factor | Supplementary | BV Ricos (minimum)
Rheumatoid factor | Supplementary | Other
Bicarbonate | Supplementary | Labquality
Bicarbonate | Supplementary | CLIA | 20 | 6.5134
Bicarbonate | Supplementary | Rili-BÄK
Bicarbonate | Supplementary | Noklus
Bicarbonate | Supplementary | RCPA | 10 | 3.1711
Bicarbonate | Supplementary | China WS/T 403-2024
Bicarbonate | Supplementary | SEKK Dmax
Bicarbonate | Supplementary | Belgium WIV-ISP
Bicarbonate | Supplementary | Turkey MoH 2026
Bicarbonate | Supplementary | BV EFLM (optimal)
Bicarbonate | Supplementary | BV EFLM (desirable)
Bicarbonate | Supplementary | BV EFLM (minimum)
Bicarbonate | Supplementary | BV Ricos (optimal)
Bicarbonate | Supplementary | BV Ricos (desirable)
Bicarbonate | Supplementary | BV Ricos (minimum)
Bicarbonate | Supplementary | Other
Ammonia | Supplementary | Labquality
Ammonia | Supplementary | CLIA
Ammonia | Supplementary | Rili-BÄK
Ammonia | Supplementary | Noklus
Ammonia | Supplementary | RCPA | 10 | 3.1713
Ammonia | Supplementary | China WS/T 403-2024
Ammonia | Supplementary | SEKK Dmax
Ammonia | Supplementary | Belgium WIV-ISP
Ammonia | Supplementary | Turkey MoH 2026
Ammonia | Supplementary | BV EFLM (optimal)
Ammonia | Supplementary | BV EFLM (desirable)
Ammonia | Supplementary | BV EFLM (minimum)
Ammonia | Supplementary | BV Ricos (optimal)
Ammonia | Supplementary | BV Ricos (desirable)
Ammonia | Supplementary | BV Ricos (minimum)
Ammonia | Supplementary | Other
Ethanol | Supplementary | Labquality
Ethanol | Supplementary | CLIA | 20 | 7.0909
Ethanol | Supplementary | Rili-BÄK | 9 | 3.1693
Ethanol | Supplementary | Noklus
Ethanol | Supplementary | RCPA | 10 | 3.5258
Ethanol | Supplementary | China WS/T 403-2024
Ethanol | Supplementary | SEKK Dmax
Ethanol | Supplementary | Belgium WIV-ISP
Ethanol | Supplementary | Turkey MoH 2026
Ethanol | Supplementary | BV EFLM (optimal)
Ethanol | Supplementary | BV EFLM (desirable)
Ethanol | Supplementary | BV EFLM (minimum)
Ethanol | Supplementary | BV Ricos (optimal)
Ethanol | Supplementary | BV Ricos (desirable)
Ethanol | Supplementary | BV Ricos (minimum)
Ethanol | Supplementary | Other
Homocysteine | Supplementary | Labquality
Homocysteine | Supplementary | CLIA
Homocysteine | Supplementary | Rili-BÄK
Homocysteine | Supplementary | Noklus
Homocysteine | Supplementary | RCPA | 10 | 3.4207
Homocysteine | Supplementary | China WS/T 403-2024
Homocysteine | Supplementary | SEKK Dmax
Homocysteine | Supplementary | Belgium WIV-ISP
Homocysteine | Supplementary | Turkey MoH 2026
Homocysteine | Supplementary | BV EFLM (optimal)
Homocysteine | Supplementary | BV EFLM (desirable)
Homocysteine | Supplementary | BV EFLM (minimum)
Homocysteine | Supplementary | BV Ricos (optimal)
Homocysteine | Supplementary | BV Ricos (desirable)
Homocysteine | Supplementary | BV Ricos (minimum)
Homocysteine | Supplementary | Other
Lithium | Supplementary | Labquality
Lithium | Supplementary | CLIA | 15 | 5.5068
Lithium | Supplementary | Rili-BÄK | 6 | 2.0874
Lithium | Supplementary | Noklus
Lithium | Supplementary | RCPA | 10 | 3.6071
Lithium | Supplementary | China WS/T 403-2024
Lithium | Supplementary | SEKK Dmax | 12 | 4.367
Lithium | Supplementary | Belgium WIV-ISP
Lithium | Supplementary | Turkey MoH 2026
Lithium | Supplementary | BV EFLM (optimal)
Lithium | Supplementary | BV EFLM (desirable)
Lithium | Supplementary | BV EFLM (minimum)
Lithium | Supplementary | BV Ricos (optimal)
Lithium | Supplementary | BV Ricos (desirable)
Lithium | Supplementary | BV Ricos (minimum)
Lithium | Supplementary | Other
Gentamicin | Supplementary | Labquality
Gentamicin | Supplementary | CLIA | 25 | 5.4378
Gentamicin | Supplementary | Rili-BÄK | 25 | 5.4378
Gentamicin | Supplementary | Noklus
Gentamicin | Supplementary | RCPA | 10 | 2.1244
Gentamicin | Supplementary | China WS/T 403-2024
Gentamicin | Supplementary | SEKK Dmax
Gentamicin | Supplementary | Belgium WIV-ISP
Gentamicin | Supplementary | Turkey MoH 2026
Gentamicin | Supplementary | BV EFLM (optimal)
Gentamicin | Supplementary | BV EFLM (desirable)
Gentamicin | Supplementary | BV EFLM (minimum)
Gentamicin | Supplementary | BV Ricos (optimal)
Gentamicin | Supplementary | BV Ricos (desirable)
Gentamicin | Supplementary | BV Ricos (minimum)
Gentamicin | Supplementary | Other
Vancomycin | Supplementary | Labquality
Vancomycin | Supplementary | CLIA | 15 | 3.7704
Vancomycin | Supplementary | Rili-BÄK | 12 | 3.0116
Vancomycin | Supplementary | Noklus
Vancomycin | Supplementary | RCPA | 10 | 2.5058
Vancomycin | Supplementary | China WS/T 403-2024
Vancomycin | Supplementary | SEKK Dmax
Vancomycin | Supplementary | Belgium WIV-ISP
Vancomycin | Supplementary | Turkey MoH 2026
Vancomycin | Supplementary | BV EFLM (optimal)
Vancomycin | Supplementary | BV EFLM (desirable)
Vancomycin | Supplementary | BV EFLM (minimum)
Vancomycin | Supplementary | BV Ricos (optimal)
Vancomycin | Supplementary | BV Ricos (desirable)
Vancomycin | Supplementary | BV Ricos (minimum)
Vancomycin | Supplementary | Other
Valproate | Supplementary | Labquality
Valproate | Supplementary | CLIA | 20 | 4.5118
Valproate | Supplementary | Rili-BÄK | 11.5 | 2.58
Valproate | Supplementary | Noklus
Valproate | Supplementary | RCPA | 10 | 2.2391
Valproate | Supplementary | China WS/T 403-2024
Valproate | Supplementary | SEKK Dmax
Valproate | Supplementary | Belgium WIV-ISP
Valproate | Supplementary | Turkey MoH 2026
Valproate | Supplementary | BV EFLM (optimal)
Valproate | Supplementary | BV EFLM (desirable)
Valproate | Supplementary | BV EFLM (minimum)
Valproate | Supplementary | BV Ricos (optimal)
Valproate | Supplementary | BV Ricos (desirable)
Valproate | Supplementary | BV Ricos (minimum)
Valproate | Supplementary | Other
Carbamazepine | Supplementary | Labquality
Carbamazepine | Supplementary | CLIA | 20 | 5.9961
Carbamazepine | Supplementary | Rili-BÄK | 12 | 3.5667
Carbamazepine | Supplementary | Noklus
Carbamazepine | Supplementary | RCPA | 10 | 2.9593
Carbamazepine | Supplementary | China WS/T 403-2024
Carbamazepine | Supplementary | SEKK Dmax
Carbamazepine | Supplementary | Belgium WIV-ISP
Carbamazepine | Supplementary | Turkey MoH 2026
Carbamazepine | Supplementary | BV EFLM (optimal)
Carbamazepine | Supplementary | BV EFLM (desirable)
Carbamazepine | Supplementary | BV EFLM (minimum)
Carbamazepine | Supplementary | BV Ricos (optimal)
Carbamazepine | Supplementary | BV Ricos (desirable)
Carbamazepine | Supplementary | BV Ricos (minimum)
Carbamazepine | Supplementary | Other
Paracetamol | Supplementary | Labquality
Paracetamol | Supplementary | CLIA | 15 | 3.5375
Paracetamol | Supplementary | Rili-BÄK
Paracetamol | Supplementary | Noklus
Paracetamol | Supplementary | RCPA | 10 | 2.3419
Paracetamol | Supplementary | China WS/T 403-2024
Paracetamol | Supplementary | SEKK Dmax
Paracetamol | Supplementary | Belgium WIV-ISP
Paracetamol | Supplementary | Turkey MoH 2026
Paracetamol | Supplementary | BV EFLM (optimal)
Paracetamol | Supplementary | BV EFLM (desirable)
Paracetamol | Supplementary | BV EFLM (minimum)
Paracetamol | Supplementary | BV Ricos (optimal)
Paracetamol | Supplementary | BV Ricos (desirable)
Paracetamol | Supplementary | BV Ricos (minimum)
Paracetamol | Supplementary | Other
Free T4 | Supplementary | Labquality
Free T4 | Supplementary | CLIA | 15 | 4.0742
Free T4 | Supplementary | Rili-BÄK | 13 | 3.5058
Free T4 | Supplementary | Noklus
Free T4 | Supplementary | RCPA | 12 | 3.2217
Free T4 | Supplementary | China WS/T 403-2024 | 20 | 5.495
Free T4 | Supplementary | SEKK Dmax | 13 | 3.5058
Free T4 | Supplementary | Belgium WIV-ISP
Free T4 | Supplementary | Turkey MoH 2026 | 20 | 5.495
Free T4 | Supplementary | BV EFLM (optimal)
Free T4 | Supplementary | BV EFLM (desirable)
Free T4 | Supplementary | BV EFLM (minimum)
Free T4 | Supplementary | BV Ricos (optimal)
Free T4 | Supplementary | BV Ricos (desirable) | 8 | 2.085
Free T4 | Supplementary | BV Ricos (minimum)
Free T4 | Supplementary | Other
Free T3 | Supplementary | Labquality
Free T3 | Supplementary | CLIA
Free T3 | Supplementary | Rili-BÄK | 13 | 2.8193
Free T3 | Supplementary | Noklus
Free T3 | Supplementary | RCPA | 20 | 4.4769
Free T3 | Supplementary | China WS/T 403-2024 | 20 | 4.4769
Free T3 | Supplementary | SEKK Dmax
Free T3 | Supplementary | Belgium WIV-ISP
Free T3 | Supplementary | Turkey MoH 2026
Free T3 | Supplementary | BV EFLM (optimal)
Free T3 | Supplementary | BV EFLM (desirable)
Free T3 | Supplementary | BV EFLM (minimum)
Free T3 | Supplementary | BV Ricos (optimal)
Free T3 | Supplementary | BV Ricos (desirable)
Free T3 | Supplementary | BV Ricos (minimum)
Free T3 | Supplementary | Other
NT-proBNP | Supplementary | Labquality
NT-proBNP | Supplementary | CLIA | 30 | 7.7283
NT-proBNP | Supplementary | Rili-BÄK | 15 | 3.7985
NT-proBNP | Supplementary | Noklus
NT-proBNP | Supplementary | RCPA | 20 | 5.1085
NT-proBNP | Supplementary | China WS/T 403-2024
NT-proBNP | Supplementary | SEKK Dmax
NT-proBNP | Supplementary | Belgium WIV-ISP
NT-proBNP | Supplementary | Turkey MoH 2026
NT-proBNP | Supplementary | BV EFLM (optimal)
NT-proBNP | Supplementary | BV EFLM (desirable)
NT-proBNP | Supplementary | BV EFLM (minimum)
NT-proBNP | Supplementary | BV Ricos (optimal)
NT-proBNP | Supplementary | BV Ricos (desirable)
NT-proBNP | Supplementary | BV Ricos (minimum)
NT-proBNP | Supplementary | Other
PTH | Supplementary | Labquality
PTH | Supplementary | CLIA | 30 | 6.073
PTH | Supplementary | Rili-BÄK
PTH | Supplementary | Noklus
PTH | Supplementary | RCPA | 12 | 2.253
PTH | Supplementary | China WS/T 403-2024 | 30 | 6.073
PTH | Supplementary | SEKK Dmax
PTH | Supplementary | Belgium WIV-ISP
PTH | Supplementary | Turkey MoH 2026
PTH | Supplementary | BV EFLM (optimal)
PTH | Supplementary | BV EFLM (desirable)
PTH | Supplementary | BV EFLM (minimum)
PTH | Supplementary | BV Ricos (optimal)
PTH | Supplementary | BV Ricos (desirable)
PTH | Supplementary | BV Ricos (minimum)
PTH | Supplementary | Other
Vitamin B12 (total) | Supplementary | Labquality
Vitamin B12 (total) | Supplementary | CLIA | 25 | 4.8554
Vitamin B12 (total) | Supplementary | Rili-BÄK | 25 | 4.8554
Vitamin B12 (total) | Supplementary | Noklus
Vitamin B12 (total) | Supplementary | RCPA | 15 | 2.5676
Vitamin B12 (total) | Supplementary | China WS/T 403-2024
Vitamin B12 (total) | Supplementary | SEKK Dmax
Vitamin B12 (total) | Supplementary | Belgium WIV-ISP
Vitamin B12 (total) | Supplementary | Turkey MoH 2026 | 25 | 4.8554
Vitamin B12 (total) | Supplementary | BV EFLM (optimal)
Vitamin B12 (total) | Supplementary | BV EFLM (desirable)
Vitamin B12 (total) | Supplementary | BV EFLM (minimum)
Vitamin B12 (total) | Supplementary | BV Ricos (optimal)
Vitamin B12 (total) | Supplementary | BV Ricos (desirable)
Vitamin B12 (total) | Supplementary | BV Ricos (minimum)
Vitamin B12 (total) | Supplementary | Other
Active B12 | Supplementary | Labquality
Active B12 | Supplementary | CLIA
Active B12 | Supplementary | Rili-BÄK
Active B12 | Supplementary | Noklus
Active B12 | Supplementary | RCPA
Active B12 | Supplementary | China WS/T 403-2024
Active B12 | Supplementary | SEKK Dmax
Active B12 | Supplementary | Belgium WIV-ISP
Active B12 | Supplementary | Turkey MoH 2026
Active B12 | Supplementary | BV EFLM (optimal)
Active B12 | Supplementary | BV EFLM (desirable)
Active B12 | Supplementary | BV EFLM (minimum)
Active B12 | Supplementary | BV Ricos (optimal)
Active B12 | Supplementary | BV Ricos (desirable)
Active B12 | Supplementary | BV Ricos (minimum)
Active B12 | Supplementary | Other
Folate | Supplementary | Labquality
Folate | Supplementary | CLIA | 30 | 2.277
Folate | Supplementary | Rili-BÄK | 25 | 1.8736
Folate | Supplementary | Noklus
Folate | Supplementary | RCPA | 25 | 1.8736
Folate | Supplementary | China WS/T 403-2024 | 30 | 2.277
Folate | Supplementary | SEKK Dmax
Folate | Supplementary | Belgium WIV-ISP
Folate | Supplementary | Turkey MoH 2026
Folate | Supplementary | BV EFLM (optimal)
Folate | Supplementary | BV EFLM (desirable)
Folate | Supplementary | BV EFLM (minimum)
Folate | Supplementary | BV Ricos (optimal)
Folate | Supplementary | BV Ricos (desirable)
Folate | Supplementary | BV Ricos (minimum)
Folate | Supplementary | Other
Procalcitonin | Supplementary | Labquality
Procalcitonin | Supplementary | CLIA
Procalcitonin | Supplementary | Rili-BÄK | 18 | 5.4063
Procalcitonin | Supplementary | Noklus
Procalcitonin | Supplementary | RCPA
Procalcitonin | Supplementary | China WS/T 403-2024
Procalcitonin | Supplementary | SEKK Dmax
Procalcitonin | Supplementary | Belgium WIV-ISP
Procalcitonin | Supplementary | Turkey MoH 2026
Procalcitonin | Supplementary | BV EFLM (optimal)
Procalcitonin | Supplementary | BV EFLM (desirable)
Procalcitonin | Supplementary | BV EFLM (minimum)
Procalcitonin | Supplementary | BV Ricos (optimal)
Procalcitonin | Supplementary | BV Ricos (desirable)
Procalcitonin | Supplementary | BV Ricos (minimum)
Procalcitonin | Supplementary | Other
Cortisol | Supplementary | Labquality
Cortisol | Supplementary | CLIA | 20 | 6.1364
Cortisol | Supplementary | Rili-BÄK | 16 | 4.8793
Cortisol | Supplementary | Noklus
Cortisol | Supplementary | RCPA | 15 | 4.5651
Cortisol | Supplementary | China WS/T 403-2024 | 20 | 6.1364
Cortisol | Supplementary | SEKK Dmax | 16 | 4.8793
Cortisol | Supplementary | Belgium WIV-ISP
Cortisol | Supplementary | Turkey MoH 2026
Cortisol | Supplementary | BV EFLM (optimal)
Cortisol | Supplementary | BV EFLM (desirable)
Cortisol | Supplementary | BV EFLM (minimum)
Cortisol | Supplementary | BV Ricos (optimal)
Cortisol | Supplementary | BV Ricos (desirable) | 22.8 | 7.0163
Cortisol | Supplementary | BV Ricos (minimum)
Cortisol | Supplementary | Other
hCG | Supplementary | Labquality
hCG | Supplementary | CLIA | 18 | 4.1162
hCG | Supplementary | Rili-BÄK | 14 | 3.1829
hCG | Supplementary | Noklus
hCG | Supplementary | RCPA | 10 | 2.2497
hCG | Supplementary | China WS/T 403-2024
hCG | Supplementary | SEKK Dmax
hCG | Supplementary | Belgium WIV-ISP
hCG | Supplementary | Turkey MoH 2026 | 30 | 6.916
hCG | Supplementary | BV EFLM (optimal)
hCG | Supplementary | BV EFLM (desirable)
hCG | Supplementary | BV EFLM (minimum)
hCG | Supplementary | BV Ricos (optimal)
hCG | Supplementary | BV Ricos (desirable)
hCG | Supplementary | BV Ricos (minimum)
hCG | Supplementary | Other
C-peptide | Supplementary | Labquality
C-peptide | Supplementary | CLIA
C-peptide | Supplementary | Rili-BÄK
C-peptide | Supplementary | Noklus
C-peptide | Supplementary | RCPA | 12 | 2.2698
C-peptide | Supplementary | China WS/T 403-2024 | 20 | 3.8628
C-peptide | Supplementary | SEKK Dmax
C-peptide | Supplementary | Belgium WIV-ISP
C-peptide | Supplementary | Turkey MoH 2026
C-peptide | Supplementary | BV EFLM (optimal)
C-peptide | Supplementary | BV EFLM (desirable)
C-peptide | Supplementary | BV EFLM (minimum)
C-peptide | Supplementary | BV Ricos (optimal)
C-peptide | Supplementary | BV Ricos (desirable)
C-peptide | Supplementary | BV Ricos (minimum)
C-peptide | Supplementary | Other
CEA | Supplementary | Labquality
CEA | Supplementary | CLIA | 15 | 2.6113
CEA | Supplementary | Rili-BÄK | 14 | 2.405
CEA | Supplementary | Noklus
CEA | Supplementary | RCPA | 12 | 1.9924
CEA | Supplementary | China WS/T 403-2024 | 25 | 4.6744
CEA | Supplementary | SEKK Dmax | 16 | 2.8176
CEA | Supplementary | Belgium WIV-ISP
CEA | Supplementary | Turkey MoH 2026
CEA | Supplementary | BV EFLM (optimal)
CEA | Supplementary | BV EFLM (desirable)
CEA | Supplementary | BV EFLM (minimum)
CEA | Supplementary | BV Ricos (optimal)
CEA | Supplementary | BV Ricos (desirable) | 24.7 | 4.6125
CEA | Supplementary | BV Ricos (minimum)
CEA | Supplementary | Other
CA 125 | Supplementary | Labquality
CA 125 | Supplementary | CLIA | 20 | 4.3507
CA 125 | Supplementary | Rili-BÄK | 16 | 3.4463
CA 125 | Supplementary | Noklus
CA 125 | Supplementary | RCPA | 12 | 2.5419
CA 125 | Supplementary | China WS/T 403-2024 | 25 | 5.4811
CA 125 | Supplementary | SEKK Dmax
CA 125 | Supplementary | Belgium WIV-ISP
CA 125 | Supplementary | Turkey MoH 2026
CA 125 | Supplementary | BV EFLM (optimal)
CA 125 | Supplementary | BV EFLM (desirable)
CA 125 | Supplementary | BV EFLM (minimum)
CA 125 | Supplementary | BV Ricos (optimal)
CA 125 | Supplementary | BV Ricos (desirable)
CA 125 | Supplementary | BV Ricos (minimum)
CA 125 | Supplementary | Other
Anti-TPO | Supplementary | Labquality
Anti-TPO | Supplementary | CLIA
Anti-TPO | Supplementary | Rili-BÄK
Anti-TPO | Supplementary | Noklus
Anti-TPO | Supplementary | RCPA
Anti-TPO | Supplementary | China WS/T 403-2024
Anti-TPO | Supplementary | SEKK Dmax
Anti-TPO | Supplementary | Belgium WIV-ISP
Anti-TPO | Supplementary | Turkey MoH 2026
Anti-TPO | Supplementary | BV EFLM (optimal)
Anti-TPO | Supplementary | BV EFLM (desirable)
Anti-TPO | Supplementary | BV EFLM (minimum)
Anti-TPO | Supplementary | BV Ricos (optimal)
Anti-TPO | Supplementary | BV Ricos (desirable)
Anti-TPO | Supplementary | BV Ricos (minimum)
Anti-TPO | Supplementary | Other
Anti-CCP | Supplementary | Labquality
Anti-CCP | Supplementary | CLIA
Anti-CCP | Supplementary | Rili-BÄK
Anti-CCP | Supplementary | Noklus
Anti-CCP | Supplementary | RCPA
Anti-CCP | Supplementary | China WS/T 403-2024
Anti-CCP | Supplementary | SEKK Dmax
Anti-CCP | Supplementary | Belgium WIV-ISP
Anti-CCP | Supplementary | Turkey MoH 2026
Anti-CCP | Supplementary | BV EFLM (optimal)
Anti-CCP | Supplementary | BV EFLM (desirable)
Anti-CCP | Supplementary | BV EFLM (minimum)
Anti-CCP | Supplementary | BV Ricos (optimal)
Anti-CCP | Supplementary | BV Ricos (desirable)
Anti-CCP | Supplementary | BV Ricos (minimum)
Anti-CCP | Supplementary | Other
Digoxin | Supplementary | Labquality
Digoxin | Supplementary | CLIA | 15 | 2.3463
Digoxin | Supplementary | Rili-BÄK
Digoxin | Supplementary | Noklus
Digoxin | Supplementary | RCPA | 10 | 1.5546
Digoxin | Supplementary | China WS/T 403-2024
Digoxin | Supplementary | SEKK Dmax | 30 | 4.7212
Digoxin | Supplementary | Belgium WIV-ISP
Digoxin | Supplementary | Turkey MoH 2026
Digoxin | Supplementary | BV EFLM (optimal)
Digoxin | Supplementary | BV EFLM (desirable)
Digoxin | Supplementary | BV EFLM (minimum)
Digoxin | Supplementary | BV Ricos (optimal)
Digoxin | Supplementary | BV Ricos (desirable)
Digoxin | Supplementary | BV Ricos (minimum)
Digoxin | Supplementary | Other

## References
Traceability of sources — the manuscript must state the source and version of every TEa
Source | Version / year | Citation (full reference) | URL | Access date | Note
Labquality | 2026 | Labquality Oy. EQA programme evaluation limits (clinical chemistry).
CLIA | 2025 (in force 11 Jul 2024; PT providers from 1 Jan 2025) | Centers for Medicare & Medicaid Services. Clinical Laboratory Improvement Amendments of 1988 (CLIA) Proficiency Testing Regulations Related to Analytes and Acceptable Performance. Final rule, Fed Regist 2022;87:41194–242 (42 CFR 493). Table as compiled by Westgard QC, "2025 CLIA Acceptance Limits for Proficiency Testing". | https://www.westgard.com/clia-and-quality-regulation-requirements/quality-requirements/2024-clia-requirements.html |  | % limit only; "or ± absolute (greater)" recorded, not converted. Sodium, potassium, calcium have absolute limits only (no %). 25-OH vitamin D, free T3, active B12, anti-TPO, anti-CCP, PCT, C-peptide, homocysteine, ammonia not listed.
Rili-BÄK | 2023 (English 2024) | Ahmad-Nejad P, et al. Revision of the "Guideline of the German Medical Association on Quality Assurance in Medical Laboratory Examinations – Rili-BAEK". J Lab Med 2024. doi:10.1515/labmed-2024-0131. German original: Dtsch Arztebl 2023, doi:10.3238/arztebl.2023.rili_baek_QS_Labor. | https://doi.org/10.1515/labmed-2024-0131 |  | Table B 1-2 a. % used = Column 3 (permissible relative deviation of the single control measurement / relative RMSD — the IQC limit), as in the haematology and coagulation studies. Column 5 (EQA) recorded for reference. Where Column 3 is banded by concentration, the band covering most control levels is used and the Note says which.
Noklus |  | Noklus. EQA quality specification for troponin T. |  |  | Used for hs-troponin T (15%).
RCPA | December 2022 | RCPAQAP. Chemical Pathology Analytical Performance Specifications. Royal College of Pathologists of Australasia Quality Assurance Programs; December 2022. |  |  | % applies above the stated concentration; below it RCPA uses an absolute limit (recorded, not converted).
China WS/T 403-2024 | 2024 | National Health Commission of the PRC. WS/T 403-2024 Analytical quality specifications for routine analytes in clinical biochemistry. Values as compiled by Westgard QC, "Consolidated comparison of chemistry performance specifications" (updated 21 Apr 2025; accessed 18 Oct 2026). Immunoassays from the companion page "Consolidated comparison of immunoassay performance specifications". | https://www.westgard.com/clia-and-quality-regulation-requirements/quality-requirements/consolidated-comparison-of-chemistry-performance-specifications.html |  | % used above the stated concentration; absolute limit below it recorded, not converted. Bicarbonate row in the source is blood-gas pCO2 (mmHg) and is not used. CRP, lipase, hs-troponin T, hCG, B12, 25-OH vitamin D, NT-proBNP, PCT and therapeutic drugs not listed.
SEKK Dmax | 2015 | SEKK (Czech EQA programme) Dmax analytical performance specifications, as compiled by Westgard QC, "A selection of SEKK-Dmax specifications". Dmax = 2.5 × long-term (2-year) EQA reproducibility CV. | https://www.westgard.com/clia-and-quality-regulation-requirements/quality-requirements/sekk-dmax.html |  | Serum values only; selection, not the full SEKK list.
Belgium WIV-ISP | 2016 | WIV-ISP / Sciensano (Belgium). EQA analytical performance specifications, clinical chemistry, as compiled by Westgard QC, "Belgium EQA performance specifications" (updated for 2016). | https://www.westgard.com/clia-and-quality-regulation-requirements/quality-requirements/belgium-eqa-goals.html |  | Goal type per analyte (state of the art / biological variability) in the TEa input sheet.
EFLM Biological Variation Database |  | Aarsand AK, et al. EFLM Biological Variation Database. | https://biologicalvariation.eu |  | CVI/CVG point estimates; 95% CIs in the Note column where given.
Ricos (BV) | 2014 | Desirable TEa from the Ricos/Westgard biological variation database (2014 update), as listed alongside SEKK Dmax by Westgard QC. Ricos C, et al. Scand J Clin Lab Invest 1999;59:491–500 (original database). |  |  | Desirable tier only, entered directly as TEa %; optimal/minimum not filled. Superseded by EFLM where EFLM data exist.
Other: RfB |  | Referenzinstitut für Bioanalytik (RfB). EQA evaluation limit for 25-OH vitamin D (30%). |  |  | Version/year and exact programme to be added.
Turkey MoH 2026 | 2026 | T.C. Sağlık Bakanlığı, Sağlık Hizmetleri Genel Müdürlüğü, Tetkik ve Teşhis Hizmetleri Daire Başkanlığı. Tıbbi Laboratuvarlar ve Genetik Hastalıklar Tanı ve Değerlendirme Merkezleri için Analitik Kalite Güvence Esasları, EK-1. Ek-1/A Tablo 1: Kantitatif Ölçümler İçin Analitik Performans Kriterleri. Ankara; 2026. |  |  | DKD sapma sınırı = EQA deviation limit (comparison TEa). IKK CV = maximum IQC CV (reference only). Absolute limits apply around the EQA sample target and are not converted. hs-troponin T and most immunoassays here are not listed.